This document provides Guidelines and Tools that product development teams can use to increase the effectiveness of their product development efforts.

This document is intended to provide guidance on business processes only; refer to QAP4.1 for design control compliance requirements.  Contents of this guidebook are not to supersede any part of Covidien Neurovascular quality system nor procedure.

Revisions to this guidebook must be approved by:

- Director of Research and Development or Higher

- Design Quality Engineering Manager or Higher

- Regulatory Affairs Manager or Higher

Revision History:

| Revision | Effective Date | Author(s) | Description of Changes |
| --- | --- | --- | --- |
| DRAFT | N/A | S2 Insight Steve Robertson Lester Oestreich | PDP v2.0 Draft |
| A | 15MAR14 | Misty Coburn Steve Robertson Lester Oestreich | PDP v3.0 |
| B | 28MAY14 | Misty Coburn Steve Robertson Lester Oestreich | Change title of VP of Operations to Senior Director Manufacturing |
| C | 07AUG14 | Debra Andrews Steve Robertson Lester Oestreich | Clarify wording on requirements to waive Clinical Development Plan, Clinical Study and/or PPT. |
| D | 23FEB16 | Debra Andrews Lester Oestreich | Align to QAP4.1 alignment to Medtronic Design Control Procedure, add marketing claims matrix artifact, change requirement for clinical trials, PPT, etc. to optional except in specific circumstances. |

[THIS PAGE INTENTIONALLY LEFT BLANK]

# To the Reader

Welcome to the Covidien Neurovascular Product Development Process Guide Version 3.0 (PDP3).

This process guide is designed to assist you in understanding the neurovascular (NV) development process.  Our goals are to encourage innovation, to make development more predictable, and to allow the product development team to focus their time on actual design and development rather than deciphering the product development process.

Using a well-defined, flexible process that fits the needs of projects of varying complexity will allow NV to better control time, cost, and development risk, and to more rapidly launch new products and product enhancements.

Another benefit of PDP3 is that it enables systematic and continual process improvement.  We ask that you follow the enclosed guidelines, but that you also question them and help us identify opportunities for further optimization and improvement.

PMI, PMP and PMBOK are registered trademarks of the Project Management Institute.

Microsoft and Microsoft Office Word, Excel, and Project are registered trademarks of Microsoft Corporation.

[THIS PAGE INTENTIONALLY LEFT BLANK]

# Table of Contents

Page

To the Reader	iii

Table of Contents	v

Introduction	7

Scope of this Guidebook	7

Document Conventions	7

PDP Objectives	7

PDP Phases	9

PDP3 Swim Lane Diagram	11

PDP3 Integration with Project Management Methodology (PMM)	13

The PDP3 RACI	19

PDP Reviews	21

PDP3 Relationship to Quality System Documentation	25

Getting Started – Project Initiation	26

Phase 1 – Concept Phase	28

Overview	28

Key Accomplishments	28

Concept Phase Description	29

Phase 1 – Concept Phase – Reviews	35

Phase 2 – Feasibility Phase	36

Overview	36

Key Accomplishments	36

Feasibility Phase Description	37

Phase 2 – Feasibility Phase – Reviews	50

Phase 3 – Development Phase	52

Overview	52

Key Accomplishments	52

Development Phase Description	53

Phase 3 – Development Phase Reviews	63

Phase 4 – Qualification Phase	64

Overview	64

Key Accomplishments	64

Qualification Phase Description	65

Phase 4 – Qualification Phase Reviews	73

Phase 5 – Launch Phase	74

Overview	74

Key Accomplishments	74

Launch Phase Description	75

Phase 5 – Launch Phase Reviews	79

PDP Tailoring	80

Project Classification	80

Tailoring Plan	83

Tailoring Examples	84

Combining Phases	85

Tailoring Phase and Design Reviews	85

Dealing with Previously Developed Products	87

Appendices	88

Appendix A – Process Considerations	90

A1 – Gathering the Voice of the Customer	92

A2 – Lean Product Development	94

A3 – Lean Terminology	96

A4 – Design Convergence of Different Concepts	100

A5 – Characteristics of Good Requirements	104

A6 – Developing a Robust Design	106

A7 – Developing a Robust Process	116

A8 – Verifying and Validating the Product Design	120

A9– Design Transfer	128

A10 –Process Validation	130

A11 – Physician Preference Testing (PPT)	136

A12 – Lifecycle Management	138

A13 – Manufacturing Transfer	140

Appendix B – Changes Before and After Launch	144

Design Changes Prior to Product Launch	144

Design Changes After Product Launch	144

Appendix C – Configuration Management and Document Revision Levels	146

Appendix D – Regulatory Requirements	148

Regulatory Submissions	148

Appendix E – Device Master Records	152

Appendix F – Deliverable Descriptions	154

Appendix G – Glossary	166

Appendix H – References	169

# Introduction

## Scope of this Guidebook

The Covidien Neurovascular PDP3 Process Guide describes the business process for developing a product from ideation to commercialization.

This document is intended to provide guidance on business processes only; refer to QAP 4.1 for design control compliance requirements.

Contents of this guidebook are not to supersede any part of Covidien Neurovascular quality system or procedures.

## Document Conventions

The pages associated with each PDP3 phase are designated by a unique color bar in the right margin to facilitate document usability.

| Chapter | Right Border Color |
| --- | --- |
| Project Initiation |  |
| Phase 1 – Concept |  |
| Phase 2 – Feasibility |  |
| Phase 3 – Development |  |
| Phase 4 – Qualification |  |
| Phase 5 – Launch |  |
| PDP Tailoring |  |

## PDP Objectives

PDP3 was designed using Lean Principles to provide a foundation for the rapid development of breakthrough and/or enhanced neurovascular products with high customer satisfaction.

With PDP3, NV project teams can “Innovate & GO!”  The considerations for this system design are discussed below:

- Innovate –

Early activities of the PDP are designed to foster innovation, with a focus on developing multiple concepts and converging those concepts with the help of our customers.  PDP3 features –

- Voice of the Customer –

  - A focus on clearly understanding the user needs and trade-offs early in the project.

  - Identification of Critical to Quality (CTQ) requirements to focus creative energy on what matters most for the customer.

- Concurrent Engineering –

  - A full cross-functional core team is engaged from the initial phase to quickly develop concepts and run them through rapid concept evaluation cycles.

- Flexible Environment –

  - Limited documentation burden – use of drafts and streamlined approvals.

  - Project teams are allowed to explore multiple concepts and technologies without prematurely committing to a final launch date.

- GO! –

As the team reaches design freeze, the PDP is designed for rapid, predictable delivery to the customer.  Though innovation is still important throughout the development effort, the project team focuses on execution of the pre-established plans to validate the design and processes and meet business objectives.  PDP3 features –

- Lean Considerations –

  - Concurrent engineering – The advanced Phase Gate system design encourages execution of activities from future phases in parallel instead of in strict series.

  - Streamlined RACI – the process for team members to approve project deliverables throughout the development cycle has been optimized to enhance speed without jeopardizing quality.

  - Short feedback loops – The use of concurrent engineering and the placement of the reviews are designed to prevent late stage product discovery (e.g., finding out that a component cannot be manufactured per the design).

  - Tailoring – In order to make sure that all types of projects (from breakthrough to design change) can GO!, a tailoring system has been developed.  The tailoring system design is focused on making sure that the activities and deliverables add value and do not create bureaucratic waste.

  - Localized Empowerment – Integrated project management and core team structures allow for faster team decision making and agility.

- Manufacturing Transfer –

  - PDP3 adds additional clarity to the mechanics of manufacturing transfer to ensure that the process and supply chain will quickly meet or exceed the business need.

  - The Launch Phase was added to ensure that customer and business needs are met and the launch is supported.

- Launch Execution –

  - Earlier and more detailed Launch Plans will help ready the organization for launch success without delays.

  - Physician Preference Testing ensures that there is strong customer acceptance and that the commercial launch will exceed expectations.

## PDP Phases

The PDP is organized into five (5) phases of development as illustrated in Figure 1 below.  The pages in the guidebook are designated by a matching color bar in the right margin to facilitate document usability.  These phase names are shared throughout Covidien.

Figure 1 – Phases of the Product Development Process (PDP)

These phases can be summarized as follows –

Phase 1 – Concept

- The product concept advances from the problem domain (user needs) to the solution domain (technical specifications).

- Major unknowns (technical or market) are examined to reduce the project and product risks.

- A number of alternative conceptual prototypes are developed and evaluated for relevance to user needs and for technical feasibility.

- Ends in concept selection.

Phase 2 – Feasibility

- The product design is optimized to ensure that it is safe, manufacturable, cost effective and reliable.

- Supply chain activities are initiated to ensure that the desired supplier quality, capacity, cost, and capability are achieved.

- The process is designed to meet the technical specifications.

- Ends in design freeze.

Phase 3 – Development

- Design verification and design validation are performed.

- The manufacturing process is finalized to ensure that target quantities, yields, and cost can be realized.

- Ends in primary design verification and validation completion.

Phase 4 – Qualification

- The manufacturing process is validated.

- Physician Preference Testing (PPT) as well as clinical trials are conducted (if required).

- The manufacturing transfer and launch strategies are developed and agreed upon.

- Ends in market launch.

Phase 5 – Launch

- Lessons learned are captured.

- Customer feedback is captured, including requests for new features as well as concerns.

- The manufacturing process for the product is further refined to improve yield, cost, capacity, etc.

- Ends in disbanding project team.

The phases, key accomplishments and deliverables are captured in this guidance document.  The term “guidance” implies a process that is adaptable to the needs of a specific project.  The intent is to provide a tool that assists cross-functional Project Teams in planning, monitoring and executing projects.

Each phase of the PDP is distinguished by its deliverables.  Each deliverable of the PDP is designed to contain the correct elements, to improve communication between functional areas, and to assure that regulatory compliance is achieved.

An easy way to see how the various functional areas contribute to the development of a new product is through a “swim lane” diagram.  The PDP Swim Lane Diagram (Figure 2) highlights the deliverables that are fundamental to each Phase.  This figure also illustrates how the various functional groups collaborate to accomplish the phase objectives.  The Swim Lane diagram shown in Figure 2 below provides a high-level view of the development process with the following information –

- Major Deliverables – Identification of key deliverables created in each phase along with the functional area primarily accountable for the deliverable.

- Information Flow – Information is developed and flows between activities as work progresses. For example, in the Concept Phase, product development starts with user needs, which are prioritized and then converted to engineering terms (technical specifications).  Multiple potential concepts are then created and evaluated against defined criteria in order to select the concept which is the best strategic fit for the customer and the business.

- Concurrent Engineering – The relative position of major activities and deliverables in each phase indicates when concurrent engineering activities should be occurring to reduce project completion time and improve results.

- Reviews & Milestones – Timing of formal reviews and milestones act as “checkpoints” within the PDP3 process. Reviews and milestones are positioned to ensure that the design is progressing as expected and that the Business Case for the product is met.

## PDP3 Swim Lane Diagram

Figure 2 – PDP3 Swim Lane Diagram

[THIS PAGE INTENTIONALLY LEFT BLANK]

## PDP3 Integration with Project Management Methodology (PMM)

PDP3 is designed to integrate with the Project Management Methodology (PMM) in order to produce a successful product.  PDP3 and the PMM are integrated through several key touch points including:

- Project Charter – The Project Charter is the agreement between the Project Sponsor / Team and executive management.  The Project Charter is described in greater detail in the Phase 1 - Concept section below.

- Project Plan – The Project Plan is the agreement between the Project Sponsor, the Project Manager, and the Project Team.  The Project defines how the project will be executed in accordance with PDP3.

- Project Team – The Project Team (Core and Extended Teams) executes the project utilizing PDP3 guidance.  Members of the Project Team will fulfill one or more of the roles specified in Table 1 and Table 2.

Table 1 – Project Team Roles

| Role | Description |
| --- | --- |
| Senior Staff | - The Senior Staff is composed of representatives from the following functions. - Senior Director Manufacturing - VP Marketing - VP Research and Development - Director Finance - Department Head Quality - Department Head Regulatory - Often included (but not required) depending on type of project / phase of development: VP Sales, Department Head Clinical |
| Project Sponsor | - Champions the project within the organization. |
| Project Manager | - Communicates project progress to management. - Creates Project Charter and Project Plan. - Schedules and facilitates Phase Gate reviews. - Assigns tasks to the core and extended teams. - Leads tailoring activities. - Facilitates timely project and technical decision making. - Ensures project complies with PDP3. - Manages project resources, schedule and budget. - Manages project risks. |
| Core Team | A typical Project Core Team is composed of the following functions – - Project Management - Design Quality Engineering - Process Development - Research & Development - Marketing NOTE – The composition of the Core Team can change based on the complexity of the project. The Core Team is responsible for – - Providing technical leadership for their functional areas. - Making technical decisions and deciding upon technical direction for the project. |
| Extended Team | - Provides technical expertise for their functional areas. A typical Project Extended Team is composed of the following functions – - The Core Team - Operations SME - Advanced Sourcing - Regulatory Affairs - Clinical Affairs - Supplier Quality Engineering |
| Functional Management | - Assigns resources to project teams. - Works with project management to resolve technical and resource issues. - Provides oversight of project team members and deliverables for their functional area. - Provides required training to project team members from their function to complete deliverables. |

Table 2 – Individual Contributor & Functional Roles & Responsibilities

| Role | Responsibilities |
| --- | --- |
| Project Management (PM) | - Project leader (see above). |
| Marketing (MKTG) | - Manages customer relationships and provides leadership on serving our customers better. - Plans and leads VOC activities. - Collects and processes VOC to create User Needs. - Creates, inputs, and updates product forecast model throughout all phases of the project. - Creates the Marketing Claims Matrix to identify the desired marketing claims for new products or new features and track progress to completion - Responsible for the generation of all original labeling content and design. - Creates and manages marketing materials and promotions. - Creates Physician Preference Test or Limited Market Release protocols and coordinates training, execution, and data collection. - Coordinates launch of the product, including interaction with the sales organization. - Creates and coordinates training activities supporting launch. - Interfaces and provides input from Reimbursement group. - Interfaces and provides input from marketing representatives in global GBU’s and RBU’s. |
| Research & Development Engineering (R&D) | - Translates User Needs in to Technical Specifications and product design. - Creates product prototypes for concept VOC activities. - Plans and conducts feasibility to create engineering specifications and converges on the final design of the product. - Develops the design concept in to a product for that meets regulations and standards for mass production and global commercialization. - Facilitates concept selection and design freeze processes. - Develops test methods to support feasibility and V&V activities. - Plans and conducts design verification and validation activities. - Assists in planning and conducting testing to support marketing claims. |
| Process Development (PD) | - Creates, acquires, and develops equipment and fixturing to support PDP build activities. - Represents the Operational internal customer on the core team. Assures product and process design meets the business case and operational metrics as defined in the Project Charter. Provides this input independently from the R&D role. - Provides product design input using Design for Xcellence (DFX) methodology. - Conducts process risk management activities. - Plans, conducts, and oversees the manufacturing of product during the PDP through process validation. - Works with Operations team member to define Manufacturing Transfer Requirements. Primary owner of the process and the builds until Manufacturing Transfer occurs. - Works with Operations team member to understand the needs of Operations and conduct impact analysis to current production lines. - Assists in supplier selection and development. - Primary owner of creation of Device Master Record (DMR) content in accordance with Operational standards. Assures proper design transfer. |
| Design Quality Engineering (DQE) | - Oversees the quality and safety of the product design and processes used to manufacture the product. - Leads and conducts risk management activities. - Leads and conducts sterilization validation activities. - Assists in the development of test methods. Typically responsible for test method qualification. - Ensures proper process controls are in place (eg, material inspections, product inspections) to mitigate risk. - Responsible for ensuring that in planning and execution that all PDP activities are conducted in compliance with internal procedures and domestic and international laws, regulations, and standards. |
| Regulatory Affairs (RA) | - Determines regulatory requirements for submission of production to domestic and international agencies. - Submits design dossier and interacts with regulatory and notified bodies. - Develops the global regulatory plan in coordination with regulatory representatives in Global GBU’s. |
| Clinical Affairs (CA) | - Ensures safe and effective use of products through assisting in the development of risk management labeling activities during the PDP. - Performs a clinical evaluation in accordance with applicable global standards. - Plans and conducts clinical trials as required. |
| Supplier Quality Engineering (SQE) | - Ensures the appropriate selection of vendors to meet product and business needs. - Development and qualification of vendors to supply required product components, assemblies and supplies. |
| Operations Subject Matter Expert (Ops SME) | - Assists in planning and review of the product and processes to ensure they meet or exceed business requirements. - Assists in planning and coordination of product builds for PDP that utilize production resources. - Keeps Operations informed of project progress and needs. - Keeps project team informed of Operational customer needs and system requirements. - Facilitates day-to-day management of production line after process validation completion. - Facilitates Manufacturing Transfer ownership exchange from process development to Operations. Assists in developing Manufacturing Transfer Plan. |
| Finance (FIN) | - Assists in the planning, budgeting, and monitoring of the PDP and product financial requirements. |
| Legal (LEG) | - Assist in the review of the product design and labeling to meet all legal requirements. - Provides intellectual property review and freedom to operate position. |
| Facilities (FAC) | - Reviews equipment, fixtures, and processes for compliance with internal business and domestic regulations and requirements concerning the environment, health and safety. |
| Independent Reviewer (IR) | - Participant in design reviews. - More than one independent reviewer may participate in a design review, but are not required. - Must be a senior member of the technical staff and should be at a minimum a Sr. Engineer and/or an Engineering Manager from an engineering function. Junior technical staff and non-technical staff should not be used as independent reviewers. - Does not have direct/indirect professional responsibility and has not participated or been involved in the product/design stage/element under review. - Ensures a fresh perspective, based on the principle that those who are too close to the design may overlook design errors. |

[THIS PAGE INTENTIONALLY LEFT BLANK]

## The PDP3 RACI

The PDP3 RACI serves the following purposes:

- Provides a project planning tool throughout all phases of PDP projects.

- Defines the scope of a project in terms of tailoring all PDP Deliverables either “Required” or “Tailored Out” of the work scope of the project.

- Provides a formal approval process for defining the level of tailoring used by the project and the resource assignments.

- Defines accountability for the completion and accuracy of each PDP phase deliverable.

The matrix includes all of the defined phase deliverables (within this guidance document), including reference procedures/guidance documents/templates, and a corresponding matrix listing which function is Responsible/Reviewer, Accountable/Owner, Consulted/Contributor, and the Informed for the completion of each deliverable.  These resources and deliverables are assessed by Senior Staff at the Concept Phase exit.

Table 3 – RACI Role Definitions

The process for completing the Tailoring Plan is provided in the PDP Tailoring section of this guidebook.

- All PDP functional groups should be involved in the development and detailed review of the Tailoring Plan.

- Utilization of the Tailoring Plan is a requirement for all PDP projects.

NOTE – Reference the PDP Tailoring section of this guidebook for more information on tailoring deliverables and activities required in PDP3 for specific needs of a project.

Figure 3 – Example of Completed Tailoring Plan (screen shot)

## PDP Reviews

Reviews are a critical part of a well-defined product development process.  The reviews defined in this PDP guidebook are used to improve the quality of technical documentation, to ensure that deliverables are complete, and to provide product and project updates.

PDP3 contains two types of reviews –

Technical – Review of work products related to the product design (e.g., test results, specification conformance matrix, etc.).

Phase Reviews - Review of work products related to business, planning, and process compliance (e.g., IP Strategy, Clinical Development Plan, Clinical Readiness, Regulatory Strategy, etc.).  These reviews must include an overview of the technical reviews and outputs held for a product.

- Each type of review can be further defined as follows:

Technical Reviews

- A technical review is a formal, comprehensive, systematic and methodical review of a design by a team whose expertise encompasses all technical aspects of the design and must include an independent reviewer(s).

- These reviews are highly cross-functional and focus on the ability to meet both customer and operations requirements.

- These reviews provide confirmation that pertinent technical data was discussed, feedback was received and that the project may proceed with the remainder of the phase or with a Phase Gate review.

Interim (Including Non-Scheduled) Phase Reviews

- A Non-Scheduled Phase Review is held when a project will move outside the bounds of its Project Charter, and does not fall in line with a scheduled Phase Review, Clinical Readiness Review or other scheduled non-technical review.

- These reviews focus on resolving specific issues that significantly impact the project and are not a complete review of all aspects of the project.

- The Core Team involves extended team members and functional management in developing and reviewing recommendations and alternatives.

- May propose changes to the project scope, resources or timing.

- Impact of the proposed alternatives need to be clearly understood.

- Impact of not resolving the issue must be presented.

- Once a decision is made, the Senior Staff must approve the revised Project Charter.

- Reestablish project deliverables, resources and timing.

- Set new bounds for functional and cross-functional reviews.

- The Senior Staff can recommend canceling the project if the issues are severe.

- Cancellation can also come as a recommendation from the team.

  - Clinical Readiness Review:

    - Also when it is the intention of a project to begin human use during the middle of a Phase, a Clinical Readiness Review is held with the same requirements as a standard Phase review.  This review goes through all of the deliverables required to be completed in order to begin first human use.  A typical Phase IV review must still be held following completion of all Phase IV activities.

Phase Reviews (Non-Technical)

- Phase Reviews are the mechanism by which the Senior Staff reviews, approves and directs projects at the end of each phase.  Senior Staff reviews the output of the team for completeness against the project plan and design control requirements.

- These reviews are event-driven and are conducted from a business perspective; they are not the same and do not replace technical reviews.

- The purpose of Phase Reviews is to assess fulfillment of the phase objectives and enable the Senior Staff to make well-informed project decisions early in the development cycle to improve time-to-market.

- Reviews progress against the Project Plan and reviews project risks.

- The outcome of each Phase Review is as follows.

- An decision to continue (GO), cancel (NO-GO), or redirect the project.

- In the special case of a redirect, the Senior Staff and Core Team agree to a clear list of issues to be addressed and a specific date when the Core Team will come back to the Senior Staff for a GO/NO-GO decision.

- In rare cases a CONDITIONAL GO may be authorized, which requires a discrete one or few deliverables to be completed prior to closing the phase, but does not require re-review of the Phase by Senior Staff.  In most cases, this should not be the team recommendation going into a Phase Review, because the team is held accountable for completeness.

- Approval decisions require an allocation of capital and human resources to successfully complete the next phase.

Table 4 – Review Classification

| Review Element | Phase Reviews (Non-Technical) | Technical Reviews |
| --- | --- | --- |
| Scheduling | The meeting is scheduled in advance (minimum 1 week notice). | The meeting is scheduled in advance (minimum 2 days’ notice). |
| Materials | Work product review material including an agenda is distributed prior to the meeting (minimum 24 hours, target 48 hours). | Work product review material including an agenda may or may not be distributed ahead of time. If not distributed ahead of time, the meeting will allow for a real-time structured review of the work product. |
| Participants | Required participants have been identified and agree to participate. | Fewer required participants have been identified and agree to participate. |
| Process | Follows a proper process and has a defined agenda. The meeting is face-to-face. | While the meeting has an agenda, the review process may only cover a specific work product. The meeting is face-to-face. |
| Leader | Facilitated by project manager. | Led by the person calling the meeting. |
| Independent Reviewer | Responsible for candid review and feedback of the content under review to the project team and staff. The independent reviewer also approves the review to ensure feedback has been appropriately considered and integrated as appropriate. | Responsible for candid review and feedback of the content under review to the project team. The independent reviewer also approves the review to ensure feedback has been appropriately considered and integrated as appropriate. |
| Meeting Minutes / Action Items | Produces a Meeting Minutes document containing findings as well as an Issues Log of issues needing resolution. | Produces a Meeting Minutes document containing findings as well as an Issues Log of issues needing resolution. |

[THIS PAGE INTENTIONALLY LEFT BLANK]

## PDP3 Relationship to Quality System Documentation

PDP3 provides guidance for development process activities which are governed under QAP 4.1 Design Controls.  Figure 4 defines how PDP3 and its accompanying documents fit into the hierarchy of the Quality System.   Additionally, for ease of use, PDP3 has been designed to work with several accompanying documents, as defined in Figure 4 below.

Figure 4 – PDP3 Document Hierarchy

As the figure shows, the following elements contribute to a complete and comprehensive product development process.

Quality Assurance Procedure QAP 4.1 – Defines the official Covidien Neurovascular design control requirements.  Compliance to QAP4.1 is mandatory per Covidien Neurovascular’s Quality System.  QAP4.1 covers a subset of the guidance materials.

PDP3 – Provides guidance on best practices for product development.

Tailoring / RACI Matrix (“PDP3 RACI”) – Microsoft Excel® document provides guidance on the content of each deliverable and required approvals.

Auxiliary Processes – Provide additional in-depth guidance related to a specific process area within the product development process.

PDP3 Templates – Provide additional guidance on the content of specific deliverables by establishing standards and examples that allow best practices to be leveraged over multiple development efforts.

[THIS PAGE INTENTIONALLY LEFT BLANK]

# Getting Started – Project Initiation

Product Development Process (PDP) projects are initiated by the NV Portfolio Management Team as part of the portfolio management business process.

Early-stage R&D or marketing exploratory activities prior to PDP initiation are sometimes referred to as “Phase 0” activities.  To more formally explore new technologies, a companion business process called the “Technology Development Process (TDP)” may be employed, by which entrance into the PDP would occur at a prescribed exit review within the TDP.  The TDP process is not covered within this Guidebook.

Entry criteria to begin Phase 1:

- Business model loaded into Enrich Software – All divisions of Covidien Vascular Therapies are required to load project financial information to calculate financial metrics, such as 10-year NPV.  This information is used by the NV Portfolio Management Team to choose and prioritize projects.

- The Portfolio Management Team approves the project

- A Project Sponsor is selected – The Portfolio Management team selects a sponsor for the project.  For new product development, this is often the Technical Director.

- The Project Sponsor selects a Project Manager.

- The Project Manager works with the Functional Managers to select a Core Team.  The team then begins working on the Project Charter (Phase 1 deliverable).

- Project Initiation Form (CF1658) is completed and approved.  This form is attached to the Project Plan in the Concept Phase.

[THIS PAGE INTENTIONALLY LEFT BLANK]

# Phase 1 – Concept Phase

## Overview

The Concept Phase starts the formal Product Development Process (PDP).  In this phase, the main goal is to evaluate the business potential and technical feasibility of one or more product concepts so that the business can make a decision whether to commit the budget, resources, and time to develop the product.

## Key Accomplishments

Gather and prioritize user needs.

Evaluate various candidate concepts and select the concept that is the best strategic fit for the business need.

Perform competitive analysis and survey the intellectual property landscape.

Identify and mitigate technology and product safety risks.

Perform detailed planning of the Feasibility Phase and high level planning of all subsequent phases for the product.

Make a Go/No-Go decision, once a viable concept is identified.

## Concept Phase Description

| Process Step | Detailed Instructions |
| --- | --- |
| - Entry into Concept Phase | The following criteria must be met to enter into the phase – Business model is loaded into Enrich Software. Project Sponsor, Project Manager, and Core Team are identified. Management has provided budget and resources to accomplish the activities of the Phase. Project Initiation form (CF1658) is completed. This form is ultimately attached to the Project Plan. |
| - Create the Project Charter | The initial activity performed in the Concept Phase is the creation of the Project Charter, which includes definition of key parameters of the product, estimated time and resources required to accomplish the scope, and the key elements of the business case that show the product is a good fit into the overall investment portfolio. Additional guidance for the Project Charter is located in DPRD01. A project team recommendation to combine PDP Phases must be justified and approved in the Project Plan [see additional details in the Tailoring Section of this Guidebook]. The Project Charter should be maintained and updated as necessary over the life of the product development effort. The identification of significant changes to the Charter signals the need for an Interim Phase Review of the proposed changes, resulting in either approval or rejection of a modified Charter and all its downstream consequences. |
| - Create the Project Plan | Another initial activity performed in the Concept Phase is the creation of the Project Plan, which includes objectives, scope, schedule, budget, and resources required for the project. Project planning is conducted per QAP4.11 with additional guidance below. Key steps in project planning include – Align on project objectives. Complete the Tailoring Plan to determine what PDP deliverables are in-scope for the project (for details see the chapter on PDP Tailoring). The Tailoring Plan is attached to the Project History File (PHF). Identify all other work products for the project. Some work products are not directly identified in the Tailoring Plan, but may be important for the project. Conduct a Kickoff Planning session involving the extended team and all critical customer & support functions to develop or validate the Project Plan including a detailed schedule (typ. Gantt). For long projects, it may be impractical to have a schedule detailed to the task level for the entire project. At a minimum, the Project Plan should provide a high level schedule for the entire project and a detailed schedule for the upcoming phase (with the exception of Phase 1 which is planned at the beginning of the phase and often will include Phase 2 detail). - The Project Plan is maintained and updated as necessary over the course of the project. |
| - Initiate the Design History File | All projects subject to design controls require a Design History File (DHF). The DHF is initiated in the Concept Phase and maintained throughout the product development cycle. The DHF shall contain or reference the records necessary to demonstrate that the design was developed in accordance with the Project Plan and QAP4.1. Key steps in initiating and maintaining the DHF include – - The DHF should be maintained as results are made available throughout each phase. - Non-mandatory “Recommended” PDP documents are not typically required to be listed in the DHF. Refer to the PDP RACI for guidance on what documents to reference in the DHF. These “Recommended” documents are typically referenced in other documents but may be added as deemed necessary by the project team. - If a document is a numeric controlled document, the DHF should be updated to include it. This document may be noted in a latter phase where the deliverable is required per the DHF Checklist template. Upon alpha release of the document, the DHF must be updated to reflect this change in the phase where the deliverable is required. - The DHF must be updated at the conclusion of each phase with deliverable result references. - The DHF must be audited by Quality Engineering and released at the conclusion of each phase to ensure accuracy, typically shortly after the Phase Review meeting. |
| - Initiate the Project History File | All project documents that are not referenced directly on the DHF are listed in the Project History File (PHF). - QAP4.1 and PDP3 RACI list documents to be referenced directly on the DHF. - All other documents to be listed in the PHF – The PHF has the deliverables listed as “required for this project” from the Tailoring Plan. PHF deliverable references specific to the completed phase are to be listed in the Phase Review slide deck, demonstrating completion of those activities. The PHF ultimately becomes part of the Phase Review records. |
| - Gain a Clear Understanding of User Needs | One of the key objectives of the Concept Phase is to gain an understanding of user needs. This assures that when released to market, the product is successful both in the eyes of customers as well as from a business perspective. Several Voice of the Customer (VOC) activities comprise this step. - Voice of the Customer Plan (VOCP) - A VOCP is approved to detail and direct the activities required to complete this process step. The VOCP is created early in this phase and is updated as needed throughout the product development cycle. - A Concept Selection Report is created towards the end of this phase and includes what has been learned from the Concept Phase activities carried out in accordance with the VOCP. - The Marketing Claims Matrix is used to identify the claims marketing desires to make regarding a new product or new feature. It outlines the supporting data and regulatory requirements needed to support the desired claims. This tool is intended to be updated throughout the development process. NOTE – Additional information on Gathering the Voice of the Customer can be found in Appendix A1. User Needs and the Specification Conformance Matrix (SCM) The User Needs are elicited through the VOC activities. Early in this phase, the most important user needs that the product must meet are documented in the first section of the Specification Conformance Matrix (SCM). The SCM is updated throughout the phase as additional needs are identified and as those needs are translated into more specific user needs. By the end of the Concept Phase a preliminary version of the SCM is reviewed. Also by the end of this phase, Critical to Quality (CTQ) elements are selected and denoted in the draft SCM. These elements represent the key requirements that must be met to satisfy the customer. NOTE – Tips for Developing Good Requirements can be found in Appendix A5. Concept Development Based on the User Needs, brainstorming of various product concepts is performed. Ideas are captured in Engineering Notebooks or preliminary rough drafts to allow prototype development to occur. Test Method Development The Design Test Method (TM) Documents are initiated in the Concept Phase to include the test methods required to evaluate the competing concepts and competitive products. Competitive Analysis Analysis of competitive products is performed and documented in the Concept Selection Report. Knowledge of competitive products can be used to help set preliminary requirements and specifications. Prototype Development - Multiple device prototypes are developed to further elicit VOC, to refine understanding of user needs, and to test the suitability of potential designs to meet requirements. Prototypes may include – Simple drawings and artwork representing product concepts. 3D models of product concepts. Actual devices produced as the result of design efforts with target materials and manufacturing processes. Input from representative customers is required to identify those requirements that are Critical to Quality (CTQ). During prototype development, the Project Team should engage the Advanced Sourcing and Supplier Quality groups for any sourcing decisions. - NOTE – Prototypes developed during the Concept Phase are always “proof of concept” and should never be considered to be the “final” design. |
| - Conduct Preliminary Engineering Design | The creation of physical prototypes requires that R&D, Process and Quality engineers begin to translate perceived user needs into technical specifications and associated design and process documentation. The key deliverables of this effort include – Preliminary Technical Specifications The Technical Specifications contain requirements that are derived from sources including – User Needs Standards and Regulatory Requirements Hazard Analysis Manufacturing Standards Technical Specifications are drafted but not fully developed at the conclusion of the Concept Phase. Technical Specifications are listed in the Specification Conformance Matrix (SCM). Multiple Device Prototypes Multiple physical prototypes are developed to support – VOC activities Feasibility studies that compare the performance of the evolving design to the requirements Risk and hazard mitigation Manufacturing planning Competing concepts are evaluated to determine the best fit for the customer and business needs. The result of the evaluation of these concepts is documented in the Concept Selection Report. NOTE – Additional information on Design Convergence of Different Concepts can be found in Appendix A4. Preliminary Design Specifications In order to produce physical prototypes, the Design Drawings will be produced. These typically are in the form of component, assembly, and top level drawings or solid models. While not “final” specifications, standard engineering practice should be followed to assure that these early drawings comply with established procedures for these documents. Product Risk Management Approval of the Risk Management Plan – Outlines the risk management activities to be performed throughout the phases of product development to ensure that the product has an acceptable level of risk. Risk Management is conducted per QAP 4.19. Approval of the Procedure Hazard Analysis (PHA), part of the Risk Management Workbook (RMW), assures that medical risks are understood early in the design. The PHA is developed in accordance with QAP 4.9. Review of Manufacturing and Assembly Processes Develop and review methods/processes used to manufacture and assemble concept models and prototypes. Develop draft manufacturing analysis, process flow model, and cost model for each prototype model planned to exit phase. Technical Risk Management Technical project risk is managed via use of project management tools like the Risk Register. - NOTE – Most deliverables in this step are draft documents and would not be considered complete or near complete. These deliverables include only the most important information required to document the major elements of the prototypes to be created. Draft documents are approved as per Appendix C. |
| - Plan Development Builds | The project manager should establish a clear line of communication with production planning and pilot line management around the development build plan and upcoming builds. A recurring meeting should be established to monitor changes in build requirements. |
| - Conduct IP Planning Session | The project manager and Legal (IP) group conduct an IP Planning Session to lay out the IP milestones and deliverables. The IP attorney signs that the session was conducted on the “IP Planning Session Record” form. Include the signed form in the Phase 1 Review slide deck. |
| - Create the Regulatory Plan | - Create and approve the Regulatory Plan to outline the regulatory strategy and approach for all intended markets. - Domestic Markets, European Markets, IDM, CAPLA, etc… should be discussed including submission strategies. - Specific regulatory requirements for FDA, CE, PMDA, CMDA, TGA, SFDA, KFDA, etc… should be discussed in detail in order to appropriately plan the feasibility, development, and testing stages of the project. - The regulatory and marketing team members should work directly with the RBU teams to integrate all international requirements in to the project via the Regulatory Plan. |
| - Draft Functional Plans | During the Concept Phase, team members begin to assess the impact of this product on their activities. Once one or more concepts that meet the business needs are identified, plans should be created for all support functions that will participate in product development. Planning for the following functions can start during the Concept Phase, but will not be completed until development of the concept is well underway (ref. PDP RACI) – Clinical Development Plan – Outlines the approach to clinical testing and potential clinical trials. – (Clinical Studies, if required to validate marketing claims or fulfill regulatory requirements, are an element that may be conducted during Design Validation.) Supplier Development Plan – Outlines the types of suppliers and components that will require assessment and qualification, and defines how suppliers will be managed. Design Characterization Matrix (DCM) – Outlines the activities and testing to be performed to develop the engineering design specifications and to define challenge testing to verify that those specifications are correct. Process Characterization Plan – Outlines the activities that will be performed to assess the feasibility and availability of different process technologies, determines parameters for qualification and production, and develops and selects the appropriate assembly flow and processes to support the manufacturing and cost of the proposed product. NOTE – Most plans are living documents and should be updated throughout the product development cycle as further knowledge is gathered. |
| - Exit from Concept Phase | The Concept Phase is exited once a Go/No-Go decision is made as a result of the Phase Gate Review (see Table 5). The following exit criteria must be met to exit this phase - Planning of the Feasibility Phase is completed. High level planning of the entire project is completed. At least one candidate concept has testing that demonstrates a good fit for both customer and business needs. The business case for the product and project is established. Remaining technical, project, and product risks are understood and mitigation plans are in place. |

## Phase 1 – Concept Phase – Reviews

Table 5 describes basic information for the reviews in the Concept Phase.  These reviews ensure that a good foundation is in place for the Feasibility Phase by ensuring that user needs are correct and complete, and that sound engineering practices have been used in evaluating and selecting the final concept(s).

Table 5 – Concept Phase – Reviews

| Review Name | Description | Inputs |
| --- | --- | --- |
| ID: R1-01 Name: User Needs Review Type: Technical | User needs are assessed to ensure the characteristics of good requirements, ensure that user needs are correctly expressed, and to identify Critical to Quality features. | Draft User Needs (in SCM) Critical to Quality Identified Approved VOCP |
| ID: R1-02 Name: Concept Selection Technical Review Type: Technical | Review of the concepts considered and the rationale behind the concept selected. | Prototype testing results Prototype VOC results |
| ID: R1-03 Name: Phase 1 Gate Review Type: Phase Review | A business review at the end of the Concept Phase to determine if the concept selected is a good fit for the customer and business needs. A Go/No-Go decision is made at the conclusion of this review. | Project Charter Project Plan All mandatory phase deliverables |

# Phase 2 – Feasibility Phase

## Overview

The main goal of the Feasibility Phase is to develop the selected concept into a design that meets clearly defined customer and product requirements, is safe and efficacious, is reliable, is projected to be manufacturable within target costs and yields, and is commercially viable.

## Key Accomplishments

Validate that the VOC has been appropriately translated into user needs and technical specifications.

Optimize the design through characterization and robustness testing.

Develop and document the manufacturing processes.  Optimize through Design for Manufacturability and Assembly, characterization and evaluation of process capabilities.

Ensure effective mitigations of product hazards.

Ensure approved suppliers exist with the capability to provide the required costs, quality, and quantities.

Prove that commercial objectives continue to be satisfied.

Make a Phase Exit decision, once feasibility and pre-verification/validation testing is completed.

## Feasibility Phase Description

| Process Step | Detailed Instructions |
| --- | --- |
| - Validate the VOC to Technical Specifications Translation | VOC activities will continue and are an essential part of the Feasibility Phase. Voice of the Customer activities in the Feasibility Phase include – Document VOC Activities – Feasibility testing is performed in the Feasibility Phase to document VOC activities. This testing is performed through consolidation of design options through an iterative process of design and testing. Functional and performance data is gathered to demonstrate that the consolidated design will, with reasonable confidence of success, meet User Needs. Baseline the User Needs – The Specification Conformance Matrix (SCM), which includes the User Needs and Technical Specifications is completed and approved. The final inputs to the User Needs come from – Competitive product analysis, within the Concept Selection Report, provides additional user needs necessary for the product to be competitive in the market. The VOC is validated with the customer as demos, prototypes, tests, data and drawings that become available from the design work being performed concurrently. Customer feedback is reviewed to determine if user needs need to be added or modified. Translation of User Needs into Technical Specifications and designs. As design work evolves and it becomes apparent that certain user needs cannot be met within the constraints of the project, the SCM is re-assessed. During this process Critical to Quality (CTQ) elements of the product and its design are determined and documented in the SCM. (Reference DPRD02.) Baseline the Technical Specifications – The Technical Specifications (also referred to as Product Specifications) are finalized and approved in the Specification Conformance Matrix (SCM). The Technical Specifications are system requirements to which the design must conform. Detailed design work ensures that the requirements are achievable given the project constraints of time, budget, resources, and technology. NOTE – The User Needs are typically base lined prior to the Technical Specifications. The Technical Specifications may evolve as Steps 2 and 3 below are conducted. Conduct Technical Specifications Review –The customer and product requirements are the foundation of all subsequent design work. To ensure that this is a solid foundation, a technical review of the requirements is performed to ensure that they possess the characteristics of good requirements and that all User Needs are accounted for in the Technical Specifications. Reference the section in Appendix 5 on Characteristics of Good Requirements. As part of the requirements reviews, Critical to Quality (CTQ) requirements are verified to ensure that they have been properly traced from the User Needs to Technical Specifications. |
| - Complete Detailed Design | The preliminary design created during the Concept Phase is evolved to ensure that all Technical Specifications can be satisfied and that the design will meet its reliability, safety, quality, cost, and performance goals. Depending on the complexity of the design, the product (or “system”) may be partitioned into two or more integrated subsystems. EXAMPLE – A system for delivering coils to an aneurysm can be decomposed into the coil, the pusher used to deliver the coils, and the detachment device. This would require design work to be performed on each of the subsystems and design outputs to be created for each of the subsystems. The partitioning of functionality to subsystems, and the interfaces between subsystems, must be carefully specified and managed to ensure integration. Key steps in completing the detailed design include – Evolve the design – Iteratively optimize the design and make design decisions based on data gathered from activities performed to assess critical areas of the design. The Design Characterization Matrix (DCM) is used to document which activities will be performed for a specific project. Activities can include – Competitive Analysis / Benchmarks Design calculations Tolerance Analysis Report Design of Experiments Design for manufacturability and assembly Finite element analysis Reliability testing Mechanical models Simulations Modeling Design test methods Design Challenge Testing Protocol(s)/Report(s) Extreme condition testing Testing to failure Assess and improve the safety of the design – Continually assess the safety of the device by performing risk management activities as defined in the Risk Management Plan (RMP), including the contents of the Risk Management Workbook (RMW). The PHA, UFMECA and DFMECA sections of the RMW are completed in accordance with QAP 4.9. The RMW is approved in the Feasibility Phase with initial risk evaluation of the following – - The Procedure Hazard Analysis (PHA) is a top-down assessment of hazards and clinical effects. - The User Failure Mode, Effects, and Criticality Analysis (UFMECA) is a sequential risk assessment used to determine the risk level associated with foreseeable misuse of the product. - The Design Failure Mode, Effects, and Criticality Analysis (DFMECA) is a bottom-up assessment of potential failures of the product and the potential effects on the end user. Create Design Outputs – Document the design through the creation of appropriate Design Drawings (in accordance with QAP 5.20 and QAP 5.2), including – Part and component drawings Assembly drawings Materials specifications Schematics Item Masters Verify requirements Verify that all of the Technical Specifications are addressed in the design. Critical to Quality (CTQ) requirements are verified to ensure that they have been properly traced from the Technical Specifications to the design outputs. NOTE – Reference Developing a Robust Design in Appendix A6 for more information on completing and optimizing the detailed design. Initiate labeling – Label development begins in this phase with the initiation of the Labels and Information for Use (IFU). Key requirements for this phase include – Identification of product packaging levels (product description, GTIN, PCN/SKU, size(s), quantity of measure). Initiation of the Product Branding Name process. Generation of Global Trade Item Number(s) (GTIN) (in accordance with LST-003). Generation of Product Catalog Number(s) (PCN) and item numbers (PCN + version). Generation of RFID product classification information with RFID service provider. Complete draft of the English version of the Instructions for Use (IFU). Requirements identified as risk controls. Requirements needed to conform to standards. Requirements identifying all the languages into which the labeling and IFU must be translated. Requirements specifying the content of all labels to appear on the packaging. Requirements for the appearance of the packaging. - Design the packaging – The packaging is designed to ensure the integrity of the device from the point of manufacture to the point of final use. Key considerations for packaging design are – - The package must keep a device sterile for the entire duration of the established shelf life throughout all the stresses and hazards generated by – Sterilization Shipping. Storage environments. - The packaged product may be subjected to – Shock and vibration. Changes in atmospheric pressure. High and low temperature extremes. High and low humidity extremes. Identify Essential Design Outputs – Design outputs that are essential to the proper functioning and safety of the device are considered essential design outputs. Those elements of the design whose failures could affect safety, effectiveness, and reliability are considered essential. Essential design outputs are identified via – Risk management activities. Design reviews. Flow down of CTQ requirements to design elements. NOTE – Reference Design for Safety in Appendix A6 for additional information on identifying essential characteristics of the product. |
| - Develop the Manufacturing Process | The manufacturing process that was developed during the Concept Phase is further refined, developed, and characterized to ensure that it can consistently manufacture products that meet reliability, safety, quality, cost, and performance objectives and specifications. Key steps in development of the manufacturing process include – Characterize the process – The Process Characterization Plan is finalized in this phase and is used to specify the activities required to assess process inputs and outputs, the range over which the process can operate without causing failures, and the capability of the process to produce the volumes required. The process characterization activities result in - Preliminary process inputs (material, equipment, tooling, software, manpower, methods). Preliminary process outputs that are measurable and provide an understanding of the impact of process variables. The development of the process parameters and process window that will be used for process qualifications, process validation, and production. Assessment of which variables have a significant impact on the outputs and which do not. Preliminary process flow diagrams. Requirements for equipment, tooling, and fixtures needed. Specification of manufacturing targets (yield, time to manufacture one unit, time resources, line layout, tact time, costs). Requirements for process control activities. Conduct an Impact Assessment (CFXXXX) to determine how the changes affect other product lines (process, materials, operators, etc.). Discuss excess and obsolete inventory impact. Requirements for capital equipment expenditures. Environmental, health and safety issues related to the manufacture of the product. - Results of the process characterization are collected throughout the Feasibility Phase and documented in the Process Characterization Report in the Development Phase. - Perform Process Failure Modes, Effects, and Criticality Analysis (PFMECA) – A Process Failure Modes, Effects, and Criticality Analysis (PFMECA), which is a section of the Risk Management Workbook (RMW), is performed, in accordance with QAP 4.18, to evaluate potential failures of the process, document the effects of the failures, and identify actions to eliminate or reduce the probability of potential failures. The PFMECA is approved in the Feasibility Phase with initial risk evaluation. The PFMECA – - Identifies potential product and process failure modes. - Assesses effects of the failures on the customer and on the next downstream process. - Identifies manufacturing and assembly process causes for the failure modes. - Identifies process variables that can be controlled to reduce the probability of the failure mode. Control the process – Based on process characterization activities and the Process Failure Mode, Effects, and Criticality Analysis (PFMECA) section of the Risk Management Workbook (RMW), the steps required to control the quality of the process outputs are identified. These steps can include – Manufacturing Inspections – Manufacturing inspections are steps added to the process for the purpose of examining, measuring, testing, or gauging that one or more characteristics of a component, part, or assembly conforms to the specified requirements for that component, part, or assembly after one or more process steps have been completed. Real-time Process Feedback Controls – Controls implemented in the process to manually or automatically monitor the output and adjust the process as necessary within the validated ranges to produce an acceptable product. Statistical Process Controls (SPC) – SPC are used to monitor a production process through the use of control charts. By collecting in-process data of the output at key steps of the production process and establishing control limits, trends in the quality of the materials or processes that may affect the quality of the product can be detected before actual failures occur. NOTE – PDP3 emphasizes a prevention approach (as opposed to “inspecting in the quality”) while performing product development tasks throughout the development phases. Early focus on robust product and process design will reduce the reliance on manufacturing or Quality Control (QC) inspections. Cost of Quality (COQ) calculations demonstrate that spending additional cost and resources in early stage development on prevention measures pays large dividends in later development phases and production. With the barrier to change increasing exponentially as the phases progress, it is imperative to challenge the design and processes early. Define Process Validation Strategy and begin validation work – During the Feasibility Phase, the Process Master Validation Matrix (PMVM) is drafted and approved at the numeric level. The PMVM defines broad requirements for all validation activities– Calibration and Maintenance Equipment Installation Qualification (IQ) Software Validation Test Method Validation Equipment (or Process ) Operational Qualification (OQ) Equipment (or Process ) Performance Qualification (PQ) Product Performance Qualification (PPQ) Process Revalidation The Process Master Validation Matrix (PMVM) is used to document and track the process validation requirements, updates to validations, subsequent validations and revalidation requirements and the output of validation activities. In the Feasibility Phase, work is started on developing and executing appropriate equipment installation and operational and performance qualifications as the design and equipment are finalized. Create Device Master Records (DMR) – A draft of the Device Master Records (DMR) is created during the Feasibility Phase. These records include, but are not limited to – Product Drawings Routers (Work Orders) Bill of Materials Packaging Specifications Labels Manufacturing Procedures Packaging and Labeling Procedures Inspection Procedures Sterilization Procedures Cleaning Methods and Specifications Equipment Requirement Specifications Equipment Drawings Preventive Maintenance Procedures Company Forms Training Requirements, Methods & Documentation Job Safety Analyses NOTE – Reference Developing a Robust Process in Appendix A7 for more information on developing, refining, and optimizing the manufacturing process design. |
| - Perform Supplier Development | Suppliers should be involved early in the development process to collaborate and utilize their ideas and to help ensure that the final design is compatible with their process capabilities. As the preliminary design of the product and the process evolves during development, identification of required components, parts, and assemblies occurs. Once identified, the Core Team in conjunction with the Advanced Sourcing group, determines whether the required part will be made in-house or purchased (make vs. buy decision). For all parts to be purchased, suppliers must be identified and qualified to provide the purchased parts. Supplier activities performed in the Feasibility Phase include – Finalize the Supplier Development Plan (SDP) – Execution of the SDP is initiated upon approval of the plan early in the Feasibility Phase. However, many supplier management activities are not completed until subsequent phases. Identify potential suppliers – One or more potential suppliers that can provide the selected parts are identified. Suppliers are selected in accordance with QAP 6.5. Risks related to the ability of the identified suppliers to supply the part are established based on – Whether the part is a custom built part or an off-the-shelf part. Complexity of its manufacture. Criticality of the part, based on effect of failure. Financial health of the supplier. Ability to provide the required quantities in the required timeframes. Overall capability of the supplier. Availability of other suppliers to provide the same part. Amount of capital investment required by the supplier. NOTE – Secondary suppliers should be established based on the amount of risk identified during the initial assessment. Add components to the Approved Supplier List (ASL) – If Covidien lot numbers are required for numeric level traceability, the Approved Supplier List (ASL) must be updated to include the applicable components as “contingent” and “pending supplier approval” in accordance with QAP 6.6 and QAP 8.4. If supplier lot numbers will be used in place of Covidien lot numbers, components do not need to be added to the ASL (until numeric release). Develop suppliers – Due to the nature of the technology used in Covidien products, there may be only a few potential suppliers for certain key materials and components. - Covidien may choose to work with suppliers that have the potential of delivering difficult-to-source materials and components to improve their quality system and capability so that they can become approved suppliers. Assess suppliers – Each potential supplier identified is assessed to determine their capability to deliver products with the desired – Quality Reliability Safety Performance Technological capability Operational capability (lead times, on-time delivery, response time) Qualify suppliers – After each candidate supplier is assessed, The Advanced Sourcing Group and Supplier Quality work to review the suppliers to ensure that they can provide the necessary quality, safety, performance, and reliability. Additional qualifications beyond the assessments can involve – Evaluating prototypes Evaluating first article(s) Supplier audits Capability and Gage Repeatability & Reproducibility studies Supplier process validations |
| - Prepare for Design Verification & Validation | In order to be ready to perform formal design verification and validation in the Development Phase, many activities need to be performed in the Feasibility Phase. The key deliverables of this effort are – Equipment Acquisition – Equipment to be used during design verification/validation must be defined, evaluated and approved prior to use in accordance with QAP 4.17. The acquisition process may be defined as follows and depends on the complexity of the equipment or fixture. - Create the Equipment Development File (EDF) for each equipment or fixture and approve the initial requirements. - Create the Equipment Requirements Specification (ERS) for the equipment. - Request quotes based on the ERS and submit capital request(s). - Order equipment and perform Factory Acceptance Test (FAT) at supplier’s facility to confirm equipment meets requirements per ERS. - Receive equipment and perform Site Acceptance Test (SAT) to confirm equipment meets requirements per ERS upon installation at Covidien facility. The SAT may be a subset of the FAT testing and may be combined with the IQ. - Perform Installation Qualification (IQ) to verify – Proper installation conditions (wiring, utilities, functionality, etc.). Calibration, preventative maintenance, and cleaning schedules. Safety features are operating correctly. User documentation, prints, drawings, and manuals are available for the equipment. The manufacturer has a recommended spare parts list. Proper environmental conditions (such as temperature and humidity) for operation. Equipment operates as per requirements. Test Method Validation – Test methods to be used during design verification and in-process testing must themselves be validated. The Design Test Method (TM) Documents are initiated in the Concept Phase to include the test methods required to evaluate the competing concepts and competitive products. In the Feasibility Phase, the Test Method Documents are finalized to include test methods required for design verification testing and a plan for qualifying these methods to ensure that the data they produce are valid. The Design Test Method Qualification Protocol(s) defines the testing to be performed to qualify test methods to be used in Design Verification and / or Design Validation. Testing performed on each protocol is summarized in a Design Test Method Qualification Report(s). The Design Test Method Capability and Suitability Summary ensures that qualified test methods used to perform the design verification or validation testing are suitable for the intended purposes and capable of producing valid results. Design Verification & Validation Planning – The planning columns of the Design Master Verification & Validation Matrix (DMVM) are completed to lay out details of the design verification and validation strategy. - Design Verification / Validation Protocols – Design Verification / Validation test protocols are drafted and/or approved as required by the DMVM. Protocols often considered in this phase are – Biocompatibility Test Protocol Design Verification Test (DVT) Protocol(s) Design Validation (DVAL) Test Protocol(s) (This may include PPT and/or clinical studies, add appropriate line to DHF if these are conducted.) Sterilization Validation Protocol Packaging Validation Protocol Chemical Testing (China FDA requirement) NOTE – Biocompatibility and Sterilization Validation Protocols may be approved and execution started during the Feasibility Phase because of the long lead times typically required to complete testing. |
| - Perform Development Builds | Product builds are performed in order to confirm various aspects of the product design and other project uses. In general, Process Development is responsible for executing and overseeing all development, validation and qualification builds (including the R&D lab, pilot line, outside vendors and/or the production floor). This includes but is not limited to – Build; Purpose. Prototypes (X# Revision Level Documentation); Concepts Prototypes Test Method Development & Validation Design Selection Feasibility Testing Design of Experiments Design characterization testing Process Characterization Market Support Testing. Manufacturing Equivalent Builds (Numeric Revision Level Documentation); Pre-DVT Design Verification Tests Design Validation Tests Installation Qualifications Operational Qualifications. Production Equivalent Builds (Alpha Revision Level Documentation); Equipment/Process Performance Qualifications Pre-PPQ PPQ Additional Design Validation Activities Physician Preference Studies Production/Clinical/Human Use Builds. |
| - Perform Pre-DVT | Prior to freezing the design, a preliminary run of verification tests is performed to assess whether the design can successfully meet verification and validation requirements. The table below compares this Pre-Design Verification Test (pre-DVT) activity to formal DVT – Pre-DVT; Formal DVT. - Can use protocols designed for formal DVT prior to their approval, or as a method of dry running the protocols. - Test procedures can be documented in engineering notebooks (though protocols are preferred).; - Uses approved protocols.. - Test results can be documented in engineering notebooks, and/or in a formal test report. - A formal test report is not required (though recommended).; - Test results are documented in the executed protocol and summarized in a test report per protocol.. - Does not require the use of statistically significant sample sizes.; - Requires statistically significant sample sizes.. - Can use Test Methods prior to validation.; - Uses validated Test Methods.. - No completion criteria. Testing to be performed will be based on areas of highest risk to passing formal testing.; - Requires execution of all testing required by the protocol.. - Designed to test numeric or redlined documentation to ensure DVT build will be successful with a minimal number of non-conformances.; - Built using numeric or alpha release documents.. |

| - Freeze the Design | Freezing the design means that the product is ready to proceed to formal design verification and validation testing. - Design Freeze Technical Review - After the design has been completed and preliminary testing indicates that the design can meet the customer and product requirements, a Design Freeze Technical Review will be held in order to obtain the team’s recommendation regarding freezing the design. - Decision to Freeze - At the conclusion of the Design Freeze Technical Review, there are two possible outcomes – - Decision to Proceed – If the design team believes that the design is capable of meeting the customer and product requirements, then the team will complete the remaining artifacts of the Feasibility Phase and recommend that a Phase Gate Review be scheduled. - Decision to Revise – If the team believes that there are still questionable aspects of the design, additional work will be planned and assigned to characterize and resolve the issues. Upon successful completion of this work, the design team will either recommend to proceed or to make further revisions, and another Design Freeze Technical Review will be scheduled if necessary. NOTE – Additional information on the Controls Before and After Design Freeze can be found in Appendix C. |
| --- | --- |
| - Complete Functional Plans | Functional plans started in the Concept Phase are updated and approved during the Feasibility Phase. In addition, the Project Plan and schedule will be updated and maintained in each phase. Plans updated and/or approved in the Feasibility Phase include - Regulatory Plan Clinical Development Plan (CDP) - If a clinical study is required per the Clinical Development Plan, a Clinical Study Protocol/Report must be included in the DHF Checklist and Tailoring Plan as a deliverable. May be waived if noted in project plan or with a memo (CF/PR) stating no CDP is necessary and approved by the same approvers. Supplier Development Plan (SDP) |
| - Develop a Preliminary Launch Plan | Launch Plan details are developed by Marketing in the Feasibility Phase and may be presented at the Phase 2 Gate Review, but the plan will not be completed until the Development Phase. The details for the Launch Plan may include – - Launch success criteria - Promotional materials needed - Launch performance metrics - Training requirements for field personnel and physicians - Roll out strategy - Strategy for Physician Preference Test (PPT) - Training plan - Refined reimbursement considerations and/or plan - Excess and Obsolescence (E&O) strategy - Quantities needed for – - PPT - Launch - Samples NOTE – Reimbursement planning is a very important aspect of project and launch planning. It is the marketing core team leaders job to seek project input from the Covidien Reimbursement group. |
| - Prepare Covidien Innovation Metrics | In the Covidien corporate system, the end of Phase 2 is the official project team “business commit.” Inputs from the project are loaded into the corporate Innovation Metrics software (“PDSystem”) and are reviewed to determine overall group effectiveness. The following information will be logged by the PMO at the conclusion of the Phase 2 review – - Phase 4 planned completion date. - First planned product sale date (if different than above). - Worldwide revenue forecast for Year 1, Year 2, and Year 3 post-launch. Preferably this is broken down per SKU, though in some cases, this may not be feasible. - This data will be compared to actual data taken at launch and subsequent years of revenue. |
| - Exit Feasibility Phase | The Feasibility Phase is exited once a design freeze decision is made as a result of the Phase 2 Gate Review (see Table 6). |

## Phase 2 – Feasibility Phase – Reviews

Table 6 describes basic information for the formal reviews in the Feasibility Phase.  These reviews ensure that a good foundation is in place for the Development Phase by ensuring that technical specifications are correct and complete, and that sound engineering practices have been used in designing the product.

Table 6 – Feasibility Phase – Reviews

| Review Name | Description | Inputs |
| --- | --- | --- |
| ID: R2-01 Name: Technical Specifications Review Type: Technical | A technical review during which the Technical Specifications are reviewed against the characteristics of good requirements. Additionally, the traceability of Technical Specifications to User Needs is reviewed. | Technical Specifications User Needs |
| ID: R2-02 Name: Design for Excellence (DFX) Review Type: Technical | A technical review during which the design is assessed for Design for Manufacturability, Design for Reliability, Design for Six Sigma, Design for Safety, and/or other relevant methodologies. | Preliminary Drawings User Needs Technical Specifications |
| ID: R2-03 Name: Design Freeze Technical Review Type: Technical | A technical review during which design outputs and Pre-DVT Results are reviewed to ensure that the design is ready to proceed to formal verification and validation. | Pre-DVT has been completed. Design Outputs have been or are ready to be base lined. |
| ID: R2-04 Name: Phase 2 Gate Review Type: Phase Review | A business review at the end of the Feasibility Phase to determine that the product design is complete and that the business case is still compelling. The decision to freeze the design will be confirmed. A Go/No-Go decision is made at the conclusion of this review. | All mandatory deliverables for the Feasibility Phase have been completed. Project Plan is up to date. DHF has been updated. |

[THIS PAGE INTENTIONALLY LEFT BLANK]

# Phase 3 – Development Phase

## Overview

The main goal of the Development Phase is to optimize the manufacturing process, from supply chain to shipping, and continue taking steps to ensure the design meets customer and product requirements and is ready to be manufactured.

## Key Accomplishments

Assess the capability of the manufacturing processes and optimize the processes to ensure they will be capable of meeting production targets and are in control.

Establish the production line to support process validations.

Ensure that supplier processes are capable and in control by refining the supply chain and finalizing component capability studies.

Verify the design meets all technical specifications.

Validate the design to ensure it meets the user needs and intended uses.

Complete the Device Master Records (DMR).

## Development Phase Description

| Process Step | Detailed Instructions |
| --- | --- |
| - Optimize the Process and Establish the Production Line | In order to complete development of the manufacturing processes, the following activities are performed – - Establish the initial production line design and equipment – Include considerations for – - Lean manufacturing, including single piece flow. - Production grade tooling. - Multiple fixtures and equipment to handle the required capacity. - Multiple-up equipment and fixturing. - Poka-yoke fixturing. - Operator safety. - Optimize the process – To ensure Equipment/Process Validation success, the process should continue to be refined and optimized. Results of this testing are documented in the Process Characterization Report and may include – - Additional limit testing and characterization. - Bringing production tooling online to eliminate or reduce additional manufacturing and/or inspection steps. - Working with the production staff to improve manufacturing and inspection procedures. - Additional process capability and throughput studies. |
| - Prepare for Process Validation | The following activities are performed in the Development Phase in preparation for Equipment/Process Validation – - Finalize the Process Master Validation Matrix (PMVM) – The PMVM is finalized and approved based on results of the PFMECA and process characterization results. - Perform Quality Planning for Production – During the Development Phase, activities are planned to ensure that the quality of the product to be manufactured can be evaluated and controlled. The PMVM provides a plan for monitoring quality activities in manufacturing. The PMVM should include – - Defined sampling plans to be used. - Description of processes that will require monitoring and actions to be performed based on monitoring results. - Any required inspections of product in preparation for or as a results of process validation should be integrated directly in to production and performed by trained production personnel. - QC inspections should not be added unless it is absolutely necessary and the reasons why production cannot perform the inspection must be justified when implemented. - Preliminary Inspection Reduction Plan to reduce the number of inspections performed in production and to move IQC inspections From Dock to Stock where feasible. The Inspection Reduction Plan defines all QC & manufacturing inspections and conditions needed to eliminate the inspection. Reference Process Validation in Appendix A10 for additional information regarding process validation. - Validate Test Methods – Test methods used in the manufacturing process must be validated. The Process Test Method Qualification Protocol(s) defines a series of experiments to demonstrate that the tests are adequate for the intended purposes. Results of process test method validation(s) are documented in the Process Test Method Qualification Report(s) and applicable manufacturing or quality procedures. Depending on the complexity of the test method being qualified, test method qualifications may be initiated in the Feasibility Phase, but completion is not required until early in the Qualification Phase. For more information related to test method validation reference the sections on Test Method Validation and Gauge Repeatability & Reproducibility in Appendix A7. - Create Equipment/Process Validation Protocols – Equipment/Process Validation Protocols include Installation Qualification (IQ), Operational Qualification (OQ), and Performance Qualification (PQ). These protocols may be created in the Feasibility Phase and are finalized early in the Development Phase. Equipment and process qualification are executed in accordance with QAP 4.17 and QAP 9.4. - Perform Equipment/Process Qualification – Equipment/Process Qualification consists of protocols and reports related to qualifying equipment used in the manufacturing process. Equipment qualification activities may be initiated in the Feasibility Phase, but completion is not required until the Development Phase. These qualifications are performed in accordance with QAP 4.17. Process qualification activities are also initiated in the Feasibility Phase, but may not be completed until early in the Qualification Phase. These qualifications are performed in accordance with QAP 9.4. Equipment/Process qualification consists of three types of qualification tests – Test; Description. Installation Qualification (IQ); - IQ verifies that the equipment has been installed correctly.. Operational Qualification (OQ); - OQ demonstrates that the process will produce acceptable results and establishes worst case limits for process parameters.. Performance Qualification (PQ); - PQ establishes long term process stability. - Stand-alone qualifications that may be completed as part of equipment development and as a precursor to PPQ.. - Elements of the IQ include – Proper installation conditions (wiring, utilities, functionality, etc.). Calibration, preventative maintenance, and cleaning schedules. Safety features are operating correctly. User documentation, prints, drawings, and manuals are available for the equipment. The manufacturer has a recommended spare parts list. Proper environmental health and safety conditions (such as temperature and humidity) for operation. Equipment operates as per requirements. - Elements of the OQ include – Process control limits (time, temperature, pressure, setup conditions, etc.). Correct setting of software parameters. The range of the raw material specifications (Multiple lots). The process operating procedures. Material handling requirements. Process change control. Operator training (multiple shifts, multiple operators). Stability and capability of the process. Possible equipment drift. Potential failure modes, action levels, and worst-case conditions based on the PFMECA. - Elements of the PQ include – Actual product and process parameters and procedures established in OQ. Acceptability of the product. Assurance of process capability as established in OQ. Process repeatability and long term process stability. Subassembly qualification when testing cannot be performed in the final assembly form. (For example, a laser weld may be built and tested as a subassembly to test the strength of the joint because this joint cannot be tested in the final assembly form.) NOTE: Individual equipment/process PQ’s may not be required if covered in the upcoming Product Performance Qualification (PPQ). - Revalidation criteria for each process are established after IQ/OQ/PQ has been performed. Revalidation of the process may be necessary if – A change(s) is made to the process that can affect the output of the process. A negative trend(s) in quality indicators is detected. The product design is changed and is determined to affect the process. The equipment is moved. The same process is used for a different purpose. - Perform Software Validation – Software validation is performed in accordance with QAP9.7. Any computer/PLC controlled equipment with software/firmware used in the manufacture or test of the device must demonstrate it meets its intended use. The major focus of software validation is to verify proper performance against pre-determined criteria applicable to the intended use of the software. Verification of proper performance can include – Software version / configuration Applicable user functions and operator controls Maximum and minimum ranges for allowed operator inputs Error and alarm conditions provided by the software Safety related functions Startup and power down sequences Security requirements (identification, password) Data retrieval and storage |
| - Initiate the Design Transfer | - NOTE – PDP3 discusses design transfer from Process Development to Operations in two steps. The first step is the transfer of the knowledge required to manufacture the device, which is called Design Transfer (initiated in this phase). The second step is the transfer of responsibility for all business-related aspects of manufacturing to Operations, which is called Manufacturing Transfer. Manufacturing Transfer is completed in the Launch Phase and is discussed in more detail in the Manufacturing Transfer section (Appendix A13). Design transfer is the introduction of a design into the production environment. Design transfer is initiated in the Development Phase to support design verification and validation builds. The purpose of the initiation of the first step is to ensure that – - Control of all the deliverables required to manufacture the product has been initiated to prepare for transfer to Operations. Transferring the Device Master Records (DMR) (including the BOM) to manufacturing is the major activity of design transfer. The draft Device Master Records (DMR) from the Feasibility Phase is further developed during the Development Phase by approving records at the numeric level revision (at a minimum). These records include, but are not limited to – Product Drawings Routers (Work Orders) Bill of Materials Labeling Specifications Packaging Specifications Labels Manufacturing Procedures Packaging and Labeling Procedures Inspection Procedures Sterilization Procedures Cleaning Methods and Specifications Equipment Requirement Specifications Equipment Drawings Preventive Maintenance Procedures Company Forms Training Requirements, Methods & Documentation NOTE – It is NOT necessary for design transfer to occur all at once. It may occur in several steps over time. Reference the section on Design Transfer in Appendix A9 for additional requirements for transferring the design from Process Development to Operations. |
| - Perform Design Verification Testing (DVT) | DVT activities include – - Complete Design Verification Protocols and/or Reports – The Design Master Verification/Validation Matrix (DMVM) approved in the Feasibility Phase will specify all protocols and reports that need to be executed in order to perform design verification. Protocols and/or reports required by the DMVM can include – - Accelerated & Real Time Aging Study Protocol - Design Verification Test (DVT) Protocol (including EO Residual testing) - Biocompatibility Test Protocol - Sterilization Validation Protocol - Packaging Validation Protocol - Chemical Testing Protocol (China FDA requirement) - Aging Builds – The Accelerated and Real Time Aging Build is the first step to executing the Accelerated and Real Time Aging Study Protocol. These builds should include a control group (T=0), the accelerated aging group(s) and the real time aging group(s), utilizing consistent materials, processes and manufacturing personnel. This practice ensures test result differences can be attributed to aging effects as opposed to other differences between the devices. - Execute Design Verification Protocols – Approved protocols are executed to provide objective evidence that design outputs conform to design inputs. Key considerations in protocol execution include – Validated Test Methods – Test methods used for verification and validation testing must be validated. For more information related to test method validation, reference the sections on Test Method Validation and Gauge Repeatability & Reproducibility in Appendix A7. Test Equipment Calibration – Test equipment used in formal design verification must be calibrated and controlled. - Complete Design Verification Reports – Reports created as a result of design verification include – - Accelerated Aging (Interim Time Point) Report - Design Verification Test (DVT) Report - Biocompatibility Test Report - Packaging Validation Report - Sterilization Validation Report - Chemical Testing Report (China FDA Requirement) - Update RMW – Once design verification testing has been completed, the PHA, UFMECA and DFMECA sections of the Risk Management Workbook (RMW) are updated with residual risks and any new information learned as a result of the testing performed. In addition, the initial risk evaluation within the PFMECA section of the RMW is updated as necessary based on any new information from the design verification testing. - Ensure test coverage for all requirements – After all protocols have been executed and testing summarized, the Specification Conformance Matrix (SCM) is updated to include the following information – - Trace each Technical Specification to at least one design verification/validation protocol. - Trace each design verification/validation protocol to a test report that summarizes the results. - Perform Design Change Assessment – The Design Change Assessment is an analysis of how design changes performed after DVT may impact the testing performed. Reference Verifying the Product Design in Appendix A8 for additional information on design verification. |
| - Perform Design Validation | Design validation can occur in parallel with design verification activities. Design validation activities include – - Complete Design Validation Protocols – The Design Master Verification/Validation Matrix (DMVM) will specify all Design Validation (DVAL) Test Protocol(s) that need to be executed in order to perform design validation. - NOTE – Design validation should be performed on devices which are manufactured using the same methods and procedures as production devices. If production devices are not used, equivalency to production devices must be demonstrated. For any differences, a justification must be provided that explains why the design validation results for the equivalent units are valid for production units. - Execute Design Validation Protocols – Approved protocols are executed to provide objective evidence that design outputs meet customer needs and requirements. - Summarize testing performed – Validation testing will be summarized in the Design Validation (DVAL) Test Report(s). - Update RMW – Once design validation testing has been completed, the PHA, UFMECA and DFMECA sections of the Risk Management Workbook (RMW) are updated with residual risks and any new information learned as a result of the testing performed. In addition, the initial risk evaluation within the PFMECA section of the RMW is updated as necessary based on any new information from the design validation testing. - Ensure appropriate test coverage for all requirements – After all protocols have been executed and testing summarized, the Specification Conformance Matrix (SCM) is updated to include the following information – - Trace each User Need to at least one design validation protocol. - Trace each design validation protocol to a test report that summarizes the results. - Perform Design Change Assessment – The Design Change Assessment (DCA) is an analysis of how design or process changes performed after design validation may impact the requirement that validation be performed on production equivalent units. Reference Validating the Product Design in Appendix A8 for additional information on design validation. |
| - Complete Product Labeling Content | In order to prepare for process validation, labeling content is finalized in the Development Phase. Labeling related activities include – Develop labeling – Label development continues in this phase by updating the Labels and Information for Use (IFU). Key requirements in this phase include – Final product branding name. Complete and approve English content of the Labels and Instructions for Use (IFU). Artwork strategy (product, packaging, labels and IFU) as well as product packaging levels (product name and description, GTIN, PCN/SKU, size(s), quantity of measure, county of origin, manufacturing address, trademark registration and patent information). Requirements identified as risk controls. Requirements needed to conform to standards. Requirements identifying all the languages into which the labeling and IFU must be translated. Requirements specifying the content of all labels to appear on the packaging. - Requirements for the appearance of the packaging. - Any additional labeling mitigations required due to – - V&V testing performed. - Updates to the RMW. - Changes to regulatory strategy. |
| - Perform Packaging Validation | - Validate the packaging process – the Packaging Validation Protocol may be created and approved in the Feasibility Phase due to long execution and testing lead times. This protocol is created to define the testing required to validate both the packaging and the packaging process. - The packaging is verified to ensure that it meets all relevant performance, safety, and regulatory requirements. - The process used to package the device must be validated. The process validation for packaging should include – Packaging equipment qualification (IQ, OQ, PQ) Final package design Verification of expiration dates on the packages Extreme ranges of the process Application and validation of any labels used as part of the process - Note: Demonstration that the packaging preserved the integrity of the product device such that the device still meets its specifications after appropriate conditioning (e.g, simulated distribution testing) is typically conducted in the Design Verification above. - Summarize the results of Packaging Validation – Once testing is completed, the Packaging Validation Report is created and approved in the Development Phase to summarize the results of all testing related to packaging. |
| - Prepare for Launch | In preparation for launch, the Launch Plan is updated with current information and is approved in the Development Phase. |
| - Perform Pre-PPQ | A preliminary test run of the Product Performance Qualification (PPQ) is initiated in the Development Phase. The pre-PPQ is used for – - Operator training - Certification of operators for critical processes - Dry run of the draft PPQ Protocol - Review of – - Manufacturing procedures - Test methods - Test forms - Quality Assurance Procedures - System Testing - Assessment of process variations, including variations due to – - Operators - Materials - Shifts Results from the Pre-PPQ are documented in the Pre-Product Performance Qualification (PPQ) Report. Pre-PPQ Results are analyzed to identify changes to the process and / or process documentation prior to the execution of the PPQ. - Note: Depending on the project plan strategy, Pre-PPQ activities may extend into the beginning of the Qualification Phase. |
| - Begin Preparing for Manufacturing Transfer | The Manufacturing Transfer Plan (MTP) identifies the requirements that must be met to successfully transfer the knowledge and responsibility for manufacturing the product from Process Development to Operations. A draft version should be created in the Development Phase with input from Operations. |
| - Complete Supplier Approvals and Component Qualifications | Supplier approvals and component qualifications are advanced and/or finalized in the Development Phase in preparation for PPQ. Typically, the components will be released with a Contingent status until completion of the PPQ and Supplier Development Report in the Qualification Phase. |
| - Exit Development Phase | Conduct a Verification & Validation Technical Review – per Table 7. Conduct a Phase 3 Gate Review – The Development Phase is exited once a Go/No-Go decision is made as a result of the Phase 3 Gate Review (see Table 7). NOTE – In general, beginning execution of the PPQ prior to completion of the Development Phase is discouraged due to the likely cost of rework. |

## Phase 3 – Development Phase Reviews

Table 7 describes basic information for the formal reviews in the Development Phase.  The Development Phase Gate Review results in a Go/No-Go decision, and is where management approves the continued funding of the project into the Qualification Phase.

Table 7 – Development Reviews

| Review Name | Description | Inputs |
| --- | --- | --- |
| ID: R3-01 Name: Verification & Validation Technical Review Type: Technical | A detailed review of verification and validation test results, including – Design verification Design validation Accelerated aging (Interim Time Point) (if complete) Real time aging (Interim Time Point) (if complete) Biocompatibility Packaging validation Sterilization validation Chemical Testing (China FDA requirement) Additionally, the following deliverables are reviewed to assess whether verification and validation activities are complete – Specification Conformance Matrix (SCM) is reviewed to ensure that all customer and product requirements can be traced to test. V&V failures are reviewed and dispositioned. Design changes implemented after V&V testing was performed are reviewed to determine the impact of the design change on the verification and validation results. RMW is reviewed to determine if any updates need to be performed based on V&V results. | Design Verification and Validation data has been summarized and known issues are documented. |
| ID: R3-02 Name: Phase 3 Gate Review Type: Phase Review | A business review at the end of the Development Phase to determine that the process design is complete and that product is considered verified and validated. A Go/No-Go decision is made at the conclusion of this review. | All deliverables for the Development Phase have been completed. DVT indicates that the design is capable of meeting the customer and product requirements. Project Plan and schedule are up to date. DHF has been updated. |

# Phase 4 – Qualification Phase

## Overview

The main goal of the Qualification Phase is to ensure that the product, processes, and company infrastructure are ready to support product launch.

## Key Accomplishments

Confirm that all critical manufacturing processes are capable of repeatedly producing products that are within specification and that meet cost and yield targets.

Ensure manufacturing is ready for full scale production.

Ensure infrastructure is in place to sell and support the product in the field.

Identify launch performance measures.

Plan to transfer the design to manufacturing.

Prepare to transfer control of manufacturing to Operations.

Submit and obtain approval from regulatory agencies.

Document that the device is safe and effective and ready for human use.

## Qualification Phase Description

| Process Step | Detailed Instructions |
| --- | --- |
| - Prepare for PPQ | Based on the Process Master Validation Matrix (PMVM), the following activities are performed in the Qualification Phase in preparation for Product Performance Qualification (PPQ) – - Finalize the Equipment/Process Qualifications – Qualification of the equipment and processes used to manufacture the product is initiated in the Feasibility or Development Phase and must be finalized in the Qualification Phase prior to PPQ. Results of equipment and process qualifications are documented in the Equipment/Process Qualification Report(s). - Finalize Pre-PPQ – Summarize the results and review as a team to determine if the process is ready for Product Performance Qualification (PPQ). The Pre-Product Performance Qualification (pre-PPQ) Report should be completed prior to the PPQ Readiness Review (or per project plan timing; prior to the PPQ report release at the latest). - Finalize Process Test Method Validations – Validation of the process test methods is initiated in the Feasibility or Development Phase and must be finalized in the Qualification Phase prior to PPQ. Results of test method validation are documented in the Process Test Method (TM) Qualification Report(s). Process test methods are documented in manufacturing and quality procedures as part of the Device Master Records (DMR). - Update Equipment Development File(s) (EDF) – In accordance with QAP 4.17, update and approve the Equipment Development File (EDF) for each qualified equipment or fixture. - Determine Continuous Process Controls – Using the results of the PFMECA (approved in the Development Phase) and results of the Equipment/Process Qualifications, controls are put in place to recognize when the equipment used in the process begins to drift. This allows for preventive maintenance of the equipment to be scheduled prior to any effect being observed on the product being produced. - Finalize the PPQ Protocol – In order to validate that the process can repeatedly manufacture product that meets specification, the final Product Performance Qualification (PPQ) Protocol is created. The PPQ Protocol defines how validation will be conducted including test parameters, final product characteristics, manufacturing equipment, and decision points on what constitutes acceptable test results. The PPQ protocol may also include requirements for Bioburden and LAL Bacterial Endotoxin testing if not handled in a separate protocol. Reference Process Validation in Appendix A10 for additional information regarding process validation. |
| - Validate the Design Transfer | - NOTE – PDP3 discusses design transfer from Process Development to Operations in two steps. The first step was initiated in the Development Phase and is the transfer of the knowledge required to manufacture the device, which is called Design Transfer. This step is completed in the Qualification Phase. The second step is the transfer of responsibility for all business-related aspects of manufacturing to Operations, which is called Manufacturing Transfer. Manufacturing Transfer is completed in the Launch Phase and is discussed in more detail in the Manufacturing Transfer section (Appendix A13). Design transfer is the introduction of a design into the production environment. Equipment/Process Qualification is performed in order to ensure that the design was correctly transferred into the Device Master Records (DMR). The purpose of this step is to ensure that – - Control of all the deliverables required to manufacture the product has been transferred to Operations. - Operations has reviewed the deliverables to ensure that they are correct and complete and contain all the information required to manufacture the product. Transferring the Device Master Records (DMR) (including the BOM, final approved Design Drawings, accurate Item Masters and final approved Labels and Information for Use (IFU)) to manufacturing using alpha-revision released documentation is the major activity of design transfer. NOTE – It is NOT necessary for design transfer to occur all at once. It may occur in several steps over time. Reference the section on Design Transfer in Appendix A9 for additional requirements for transferring the design from Process Development to Operations. - Finalize Labels and Information for Use (IFU) – Final labels are ordered so that they can be used during the PPT and launch builds. Final labels include labels on both external and internal packaging. - Complete Information for Use (IFU) and Order IFUs– The English version of accompanying literature is translated into the multiple languages required for the product and put in the final format. Final IFU conforms to standards for IFU layout and formatting. Final IFUs are ordered so they can be used during PPT and launch builds. |
| - Perform PPQ | Product Performance Qualification (PPQ) is a major part of process validation. The goal of PPQ is to establish that the process can consistently produce product that is within the defined specifications under well-defined process conditions. The following key activities are part of process validation and PPQ execution – - Perform the Product Performance Qualification (PPQ) Readiness Review – A PPQ Readiness Technical Review is performed prior to the start of the PPQ to verify that all documents, information, materials, resources, and equipment needed for PPQ are available. The PPQ Readiness Review – - Assesses test methods to ensure that they are suitable and validated. - Reviews results of Software Validation, IQ, OQ, and PQ to ensure that all results are acceptable. - Verifies that all operators that will take part in the PPQ have been trained and / or certified. - Verifies that process validation protocols are complete and approved. - Ensures that materials to be used in the PPQ builds are available and have been procured from approved suppliers per the Supplier Development Plan. - Ensures that all issues identified in Process Characterization, Equipment Qualification, Pre-PPQ or PFMECA have been addressed. - Verifies that all DMR documentation is ready for the builds. - Verifies that EDF files are appropriately updated. - Execute the PPQ Protocol – Once approval to start the PPQ has been granted by the readiness review, PPQ execution can begin. All PPQ work is performed by trained and / or certified production operators, using top level and subassembly production work orders and using design-transferred documentation, including the same methods and procedures intended for ongoing production. When the PPQ has completed execution, results are summarized in the Product Performance Qualification (PPQ) Report. - Update the DMVM, PMVM, & RMW – Once testing is completed, the Design Master Verification/Validation Matrix (DMVM) and the Process Master Validation Matrix (PMVM) are updated to add traceability to test results. Additionally, the Risk Management Workbook (RMW) (PHA, UFMECA, DFMECA, PFMECA) is updated with residual risks and to reflect any new knowledge obtained through the execution of the process validation activities. Reference the section on Process Validation in Appendix A10 for additional details and consideration. |
| - Complete Supplier Development Report | In the Qualification Phase, the Supplier Development Report is completed and should include – - Provide the list of approved suppliers. - Identify components, parts, or services where a second source supplier should be pursued. - Reference assessment results for supplier manufacturing capabilities. - Identify risks in sourcing or supplier quality. - Reference or include a Dock to Stock component plan. - Upon release of the Supplier Development Report all suppliers/components used in the manufacturing of the product will be Approved status (a requirement for First in Human use). |
| - Update the Launch Strategy | The Launch Plan is updated in the Qualification Phase to reflect any changes in the launch strategy. Possible updates include– - Launch success criteria - Promotional materials needed - Launch performance metrics and recommended duration of Launch Phase 5 - Training requirements for field personnel and physicians - Roll out strategy - Strategy for Physician Preference Test (PPT) - Quantities needed for – - PPT - Launch - Samples - Training - 3-month component supply to support manufacturing - Updates to the 3 year post-launch Forecast |
| - Perform Quality Planning for Production | During the Development Phase, Quality began planning the activities necessary to ensure that the quality of the product to be manufactured can be evaluated and controlled. Quality planning activities in the Qualification Phase include updating the Process Master Validation Matrix (PMVM) with an Inspection Reduction Plan (which may include dock to stock for IQC). |
| - Prepare to Market the Product | - Develop the Marketing Message – In preparation for selling the product, Marketing creates the message that the marketing communication process will deliver. This includes – - Determining the sales message that will he used in all marketing materials. - Identifying potential buyers and the value proposition for each. - Creating a positioning statement for the product, both within the company’s other products and with competitive products. - Develop Marketing Communications – Marketing communications are developed and selected based on the most effective means to deliver the intended message. There are many communications media that Marketing may use to communicate the sales message. The set of all such tools and messages selected to promote the product constitutes the Marketing Communications (MARCOMM) deliverable. Examples include – Advertisements Trade Shows Videos Brochures Posters Samples Models - Marketing Support Materials – R&D may need to help develop some of the Marketing support materials, such as product samples, images, demonstrations, and models. - Verify Marketing Claims – Claims made in marketing materials must be verified to be truthful. The Marketing Support Materials & Claims Testing includes a protocol and report that documents all testing that R&D performs to verify the marketing claims. Marketing claims must be established in accordance with QAP 5.14. |
| - Verify Readiness for Human Use | Prior to the device being used on humans, it must be determined that the device is safe and effective and in all ways ready for such use. The Human Use Decision is made at the Clinical Readiness Review, and is based on a review of the Design History File with a focus on product, user, and patient safety. Key elements that must be completed prior to the Human Use Decision include – - Regulatory Approvals – Regulatory Approval(s) must be available prior to the first human use of the device. If a Clinical Study is being performed, then IRB approval and IDE submission are required prior to the start of the Clinical Study. Reference Appendix D – Regulatory Requirements for additional information related to regulatory approvals. - IP Review – The responsible IP attorney must complete and sign the “IP Review Record for First in Human Readiness” form. Include the signed form in the Clinical Readiness Review slide deck. This form is only to be used if using the product in humans prior to Phase 4 closure. Note: Prior to Phase 4 closure, an additional form “IP Review Record for Launch” must also be completed. - Verification & Validation (V&V) – V&V must be completed, including – - Accelerated Aging Study (Final Time Point) Report(s) - Verifying the shelf life for which the device will be labeled. - Bioburden & LAL Bacterial Endotoxin Report - Confirming the adequacy of sterilization procedures. - Design Change Assessment - Providing documented rationale that devices used for design validation are equivalent to initial production devices. - NOTE – All discrepancies identified during design V&V must be documented and adequately addressed or resolved. - Labels & Information for Use – Final Labels and Information for Use (IFU) must be available. Labels and Information for Use (IFU) should be reviewed to ensure that no changes are required based on validation results or regulatory approvals. - Design Transfer – Process validation and supplier qualification activities are complete and documented. Design transfer is complete. Risk Management– In accordance with QAP 4.19, risk management activities must be completed and documented in the Risk Management Report (RMR) – pre-Clinical/PPT and must include a discussion of the acceptance of residual risks. The Risk Management File Index (RMFI) must also be updated to provide a set of records and other documents that are produced by risk management. - Clinical Evaluation – The Clinical Development Plan (CDP), if required, is updated in the Qualification Phase to reflect any changes in the clinical strategy and defines the clinical evaluations to be performed for a product, if any. A clinical evaluation is the assessment and analysis of clinical data pertaining to a medical device in order to verify the clinical safety and performance of the device. May be waived if noted in project plan or with a memo (CF/PR) stating no CDP is necessary and approved by the same approvers. Clinical data can be generated from - - New clinical investigations - Literature reviews of published clinical investigations - Unpublished reports from the manufacturer, including – Biological safety files Bench testing files Complaint files Other internal information The Clinical Study Protocol/Report documents the requirements and results of any clinical studies performed per the Clinical Development Plan. Only completed if clinical study/PPT is being planned. The Clinical Evaluation Report documents the results of the clinical evaluation to date, including – - Description of the device - Intended uses, indications for use, and claims - Sources for articles to be reviewed - Selection criteria for articles to be reviewed - Data analysis - Conclusions derived from the data review - NOTE – The Clinical Evaluation Report is a requirement for the CE Mark. Reference MEDDEV.2.7.1 Rev.3 Clinical Evaluation: A Guide for Manufacturers and Notified Bodies for additional details on Clinical Evaluations. |
| - Perform Physician Preference Test (PPT) Study PPT, if required to validate marketing claims or fulfill regulatory requirements, is an element that may be included as part of Design Validation in phase 4. If PPT is conducted complete these steps and update the DHF checklist appropriately. | When included as part of the project activities, the Physician Preference Test (PPT) Study is typically the first human use of the product and is performed to confirm the results of design validation testing, confirming that the product satisfies the user needs and can be used as defined in the intended uses. - Create PPT Protocol – The PPT Protocol is created to define – - Physicians and sites that will participate in the study. - A mechanism for obtaining customer feedback. - Number of procedures to be performed as part of study. - Success criteria for the PPT study. - Intended uses that will be covered as part of the study. Reference Physician Preference Testing section in Appendix A11 for additional details on the design of PPT studies. - Execute the PPT Protocol – Execution of the PPT studies can only be performed after the following activities are completed – - Regulatory approvals have been obtained. - The Clinical Readiness Review has been performed and the product is ready for human use. - The PPT Protocol is approved. - Sites and physicians participating in the study have been identified and are ready to participate. Once execution is completed, results of the study are summarized in the PPT Report. |
| - Update Risk Management | Upon completion of any Phase 4 human-use activities, the Risk Management Report is updated. In accordance with QAP 4.19, risk management activities must be completed and documented in the Risk Management Report (RMR) and must include a discussion of the acceptance of residual risks based on results of human use studies. |
| - Prepare for Manufacturing Transfer | Key activities performed in the Qualification Phase to prepare for the transfer of manufacturing responsibility to Operations are – - Planning the transfer –The Manufacturing Transfer Plan (MTP) identifies the requirements that must be met to successfully transfer the knowledge and responsibility for manufacturing the product from Process Development to Operations. - Plan the ramp up strategy – The Manufacturing Transfer Plan (MTP) specifies the number of units that are required to be built through the completion of the Launch Phase. This includes – - PPT Build - Limited launch builds - Full launch builds - Qualify essential secondary suppliers – The Manufacturing Transfer Plan (MTP) may also discuss qualification of secondary suppliers to mitigate any risk with sole-sourced components or assemblies. - Verify readiness to transfer – The Manufacturing Transfer Technical Review is performed to assess performance against the established metrics for transferring manufacturing responsibility to Operations, including – - Target costs - Target yields - Capacity - In preparation for product launch, the Advanced Sourcing group, with support from the Project Team, prepares to transfer purchasing and planning responsibilities to the Production buyer/planner. - Order materials – Launch quantity plus 3-month supply of materials is ordered to support quantities per the Launch Plan. - Transfer materials to inventory – For all materials ordered as memo-type purchase orders, engineering memo(s) are completed to transfer materials to production inventory. - Finalize standard cost – Final component cost is established with the suppliers based on the forecast in the Launch Plan. - A summary of outstanding purchase orders (suppliers, cost, expected due dates, quantity on hand, etc.) is provided to the appropriate production buyer/planner. Reference the section on Manufacturing Transfer in Appendix A13 for additional requirements for transferring responsibility to manufacturing. |
| - Train Field Personnel | Prior to launching the product, the sales force is trained to sell the product and physicians are trained on how to use the product safely. In order to accomplish this – - Based on the Launch Plan, training materials are created to support the Field / Physicians training. - The Field / Physician Training Materials are then used to train the sales force and physicians. - The Field / Physicians Training Records provide a record of all physicians and field personnel that have successfully completed training. |
| - Exit Qualification Phase | The Qualification Phase is exited once a Go/No-Go decision is made as a result of the Phase 4 Gate Review (see. Table 8). |

## Phase 4 – Qualification Phase Reviews

Table 8 describes basic information for the reviews in the Qualification Phase.  The Qualification Phase Gate Review makes the Launch Commit decision, and is where management approves the continued funding of the project into the Launch Phase.

Table 8 – Qualification Reviews

| Review Name | Description | Inputs |
| --- | --- | --- |
| ID: R4-01 Name: Product Performance Qualification (PPQ) Readiness Technical Review Type: Technical | The purpose of the review is to verify that all documents, information, materials, resources, and equipment needed for process validation are available. | Pre-PPQ results indicate that processes are capable and in control. |
| ID: R4-02 Name: Clinical Readiness Review Type: Technical | The purpose of this review is to ascertain that the product is ready for Human Use. | Design Verification and Design Validation have been completed. All required human use deliverables are complete. |
| ID: R4-03 Name: Manufacturing Transfer Technical Review Type: Technical | The purpose of the review is to verify that device design, its components, and its configuration have been correctly translated into production specifications and that Product Development is ready to begin transferring responsibility to manufacturing. A detailed review of process validation test results, including – - Equipment/Process Qualification (IQ, OQ, PQ) - Product Performance Qualification (PPQ) - Final Time Point Accelerated Aging - Real Time Aging (Interim Time Point) (if complete) A review of current operational metrics, such as cost, yield, and capacity. | The design has been frozen. All design outputs have been created. PPQ Report complete Supplier Development Report complete |
| ID: R4-04 Name: Phase Gate 4 Review Type: Phase Review | A business review at the end of the Qualification Phase to determine if the product is ready to enter into full scale production. A Go/No-Go decision is made at the conclusion of this review. | All deliverables for the Qualification Phase have been completed. Project Plan and schedule are up to date. DHF has been updated. |

# Phase 5 – Launch Phase

## Overview

The main goals of the Launch Phase are to launch the product and to complete all activities required for general release of the product.

## Key Accomplishments

Launch the product.

Review lessons learned.

Ensure manufacturing processes are stable and are operating within established parameters for yield, and cost, and capacity.

Monitor field performance and take corrective action as required.

Continue to implement the regulatory strategy and obtaining regulatory approval to sell the product in the desired countries.

Transfer independent control and responsibility for ongoing production to Operations.

Close out the project and disband the project team.

## Launch Phase Description

| - Transfer Responsibility for Manufacturing to Operations | Once the product has been launched, and Process Development has met the transfer requirements specified in the Manufacturing Transfer Plan (MTP), then the transfer of responsibility from Process Development to Operations can be completed. The following are the key transfer activities – Execute the Manufacturing Transfer Plan – Complete deliverables required per the plan. Complete any secondary supplier qualification activities as outlined by the MTP or Secondary Supplier Qualification Plan. Complete inspection reduction activities as outlined per the MTP or PMVM. Continue to update the PFMECA and RMW as needed. Complete Additional Accelerated & Real-Time Aging Study Report(s) per the Accelerated & Real-Time Aging Study Protocol. Monitor SPC and process capabilities to ensure processes are stable and operating as designed - Perform the Manufacturing Transfer Review – The Manufacturing Transfer Update Review is the point at which Operations accepts the transfer of responsibility from Process Development. This review can cover – - Current state of manufacturing - How each requirement specified in the Manufacturing Transfer Plan (MTP) for transferring responsibility to Operations was met - Secondary supplier approvals or Post-Phase 5 plans as applicable - Availability of raw materials - PFMECA recommendations - Process capability - When acceptance occurs, Operations becomes responsible for continuous improvements of the manufacturing processes. Operations can decline to accept the transfer until all requirements specified in the MTP have been met. If Operations rejects the transfer, then Process Development must either continue to address the requirements specified in the MTP or seek to modify the MTP or Charter. Refer to the Manufacturing Transfer section in Appendix A13 for additional information on the manufacturing transfer step. |
| --- | --- |
| - Monitor Launch Performance Metrics | The Launch Plan identifies the launch performance metrics that will be monitored once the product is launched, as well as the time period during which the monitoring will be performed. Launch performance metrics may be measured using – - Limited Market Release (LMR) – Restricted Distribution – Requires LMR protocol & report. - Controlled Inventory Release (CIR) – Full Release – Customer Service Limits Sales to Specific Accounts – No protocol and report are required. - Product / Market Surveillance - NOTE – The monitoring time period ensures that the Launch Phase lasts for only the pre-determined amount of time specified in the Launch Plan. Launch Performance Metrics – Launch performance metrics are selected to measure and confirm customer satisfaction. Typical launch performance metrics can include – Adverse events and Medical Device Reports (MDRs) Complaints Customer satisfaction surveys Number of units sold Market Surveillance – Conduct review of the competitive landscape, literature review, review of symposia, etc… discussing the product. Monitoring Time Period – The monitoring time period is intended to provide a definitive end to the Launch Phase. At the end of the monitoring period, if the launch performance metrics are below expectation, then corrective actions can be performed to address the identified issues. |
| - Monitor Field Performance | Once the product has been launched, field performance is monitored through the use of – Product/Market Surveillance, including – Complaints Adverse Event Reporting - Corrective Action / Preventative Action (CAPA) Processes Complaints are analyzed using the Complaint Evaluation Standard (CES), which provides a template for evaluating complaints. The CES provides – Procedure to follow when evaluating a complaint to gather maximum data Complaint codes to use when logging complaints Guidelines for recording identifying information for the product |
| - Continue the Regulatory Strategy | In the Launch Phase and beyond, Regulatory continues to execute the regulatory strategy established in the Regulatory Plan. The Regulatory Plan is updated and the Summary Technical Document (STeD) is created and approved in the Launch Phase and includes – Any changes to the regulatory strategy. Regulatory approvals by country, including any restrictions. Post approval commitments made to regulatory agencies. Strategy and requirements, including resources for post-project closure submissions. |
| - Close Out the Project | Project Closeout involves the following activities – Celebrate Success – The accomplishments of the team should be celebrated to – Recognize achievement. Foster teamwork. Build relationships. Identify Lessons Learned – The Lessons Learned Review provides an opportunity to learn from the experiences of the project and improve performance on future projects. In order for the Lessons Learned Review to be effective, the following should be considered – Lessons can include opportunities for improvement as well as institutionalizing areas of success. Force rank the opportunities for improvement. Consider using an independent facilitator to encourage open discussion and full team participation. Discuss possible ways to improve on each lesson learned. For long projects, identify lessons learned at the end of a phase instead of the end of the project. Provide the lessons learned review notes to the Project Management Office for sharing with future teams and addressing of systemic issues. Archive Project Artifacts – Ensure that all project artifacts are archived and can be accessed for future reference. Disband the Core Team – At the end of the Launch Phase, the Project Core Team can disband. |
| - Exit Launch Phase | The Launch Phase is exited and the project is complete once the following criteria are met – - Transfer to manufacturing is complete. - Time for monitoring of field performance satisfaction has elapsed and customer satisfaction is confirmed. - The Project Close-Out Review is completed and approved. |
| - Continue the Product Lifecycle | Even after the Launch Phase is complete, the organization must continue the product lifecycle to ensure product success. Manage the Product Lifecycle – These are all the activities that go into designing, manufacturing, selling, distributing, using, and disposing of a product. The Lifecycle Management information in Appendix A12 includes several key considerations for phases of the lifecycle after commercialization of the product. Manage Design and Process Changes – After the Launch Phase, design and process changes may be required as a result of – Continuous improvement Obsolescence Field issues New regulatory requirements New user needs Cost reduction Ease of manufacturing Reference Appendix B – Changes Before and After Launch for more information on the recommended approach for design and process changes. |

## Phase 5 – Launch Phase Reviews

Table 9 describes basic information for the formal reviews in the Launch Phase.  The Launch Phase Gate Review closes the project and disbands the Project Core Team.

Table 9 – Launch Phase Reviews

| Review Name | Description | Inputs |
| --- | --- | --- |
| ID: R5-01 Name: Lessons Learned Technical Review Type: Technical | The purpose of the Lessons Learned Review is to gather the Project Team (including Core and selected Extended Team members) to review the entire project from initiation to closure, and to determine what areas of the project and processes worked well and what areas did not. | Product has been launched. |
| ID: R5-02 Name: Manufacturing Transfer Update Review Type: Technical | The purpose of the review is to officially transfer the responsibility for manufacturing of the product to Operations. | Process validation has been completed. Draft manufacturing transfer data. Additional Accelerated and Real-Time aging results are available. |
| ID: R5-03 Name: Phase 5 Gate Review (Project Close-out) Type: Phase Review | The Phase 5 Gate (or Project Close-out) Review at the end of Launch Phase is the formal end of the project. After this meeting the project team can be disbanded. | All PDP deliverables for the project have been completed. Project Plan and schedule are up to date. DHF has been updated. |

# PDP Tailoring

PDP3 is flexible and can be adjusted to meet the needs of different types of product development efforts by tailoring the deliverables and activities required in PDP3 for the specific needs of the project. Tailoring ensures that only essential activities are planned and executed in keeping with the “lean PDP” philosophy on which PDP3 was built.

Tailoring is an activity that the project core team must perform at the start of the project.  The processes for tailoring the PDP are discussed in the following sections –

Project Classification

PDP Tailoring

Special Tailoring Considerations

## Project Classification

Within PDP3, product development projects are differentiated based on expected effort, cycle time and cost.  To assist the product development team with tailoring the PDP, each project will be classified based on various characteristics of the proposed product.  This classification should be completed and documented in the Project Charter.

Projects can be classified into one of four types based on complexity –

Breakthrough – Projects that involve creating the first generation of an entirely new product and / or process.  They may include one or more of the following –

- Existing therapies using a new medical technology.

- Technologies which are not a core competency of the company.

- Products for which there is no predicate device.

- New therapies.

Next Generation (Next Gen) – Projects that create new products based on a current technology or product.  They may include one or more of the following –

- Application of existing medical technology.

- Technologies which may or may not be a core competency of the company.

- Availability of predicate devices.

- Significant new features or improvements.

Derivative – Projects that refine or improve on existing product platforms to deliver value to the customer.  They can include one or more of the following –

- Application of existing medical technology.

- Technologies which are a core competency of the company.

- Availability of predicate devices.

- Expansions or evolution of existing designs (product and / or processes).  Classic “line extensions.”

- New functionality that is visible to the customer.

Design Change – Projects that refine or improve on an existing design or process, without providing any visible benefit to the customer.  They can include one or more of the following –

- Improvements in the form, fit, or function of a device.

- Labeling changes.

- Reliability improvements.

- Manufacturing process improvements.

The following criteria will be used to assist in the classification of each project, based on the characteristics of the product being designed –

Table 10 – Project Characteristics

| Characteristic | Description |
| --- | --- |
| Financial Risk | High – A large investment is required to develop the product. Medium – A moderate investment is required to develop the product. Low – A small investment is required to develop the product. |
| Clinical Risk | High – Involves a new medical technology or therapy which is breaking new ground, typically requiring clinical trials. Medium – Involves a medical technology or therapy for which a predicate device is available, but which is new to the company. Low – Involves a medical technology or therapy which is familiar to the company. |
| Regulatory Risk | High – Requires a new PMA submission. Medium – A change to an existing PMA or new 501(k) with some regulatory planning complexity. Low – Requires relatively straightforward regulatory activities (e.g. letter to file or 510(k) with clear Covidien predicate devices). |
| Market Risk | High – The product will be used to enter a market in which the company is not currently competing and which may require extensive VOC work to understand the user needs. Medium – The company understands the market, but new product features will require VOC activities to ensure customer satisfaction. Low – The company has a good understanding of the market and the product line, and will only require limited VOC activity. |
| Technology Risk | High – Requires the use of technology or materials that are new to the company, or uses the technology or materials in a significant new application. Medium – Does not require new technology or materials, but uses the technology or materials in a new application. Low – Uses existing technologies and materials in a well understood application. |
| Technical Complexity | High – Requires the development of new processes and core competencies that the company does not currently have. Medium – Does not require development of new processes or core competencies, but requires a high level of design work to be performed. Low – Requires a relatively low level of design work. |

Table 11 classifies projects based on rating against these criteria. Where the classification is unclear or ambiguous, the more conservative (higher risk) project classification should be considered.  For example, if a project has four “High” and two “Medium” level risks it should be considered a Breakthrough.  For most projects, it is typically clear from the beginning what project classification is most fitting.  This matrix helps clarify project classification when the team is unclear.  It is intentionally not fully prescriptive, to provide adaptability to the project manager and core team to choose the most fitting classification.

Table 11 – Project Classification

| PROJECT CLASSIFICATION | Financial Risk | Clinical Risk | Regulatory Risk | Market Risk | Technology Risks | Technical Complexity (Design & Process) |
| --- | --- | --- | --- | --- | --- | --- |
| Breakthrough | High | High | High | High | High | High |
| Next Generation | Medium | Medium | Medium | Medium | Medium | Medium |
| Derivative | Medium | Low | Low | Low | Low | Low |
| Design Change | Low | Low | Low | Low | Low | Medium |

## Tailoring Plan

The table below describes the steps to be followed in order to tailor the PDP based on the complexity of the project –

| Process Step | Detailed Instructions |
| --- | --- |
| - Determine Project Classification | Review Table 11 to determine the classification for the project. Project classification is recommended by the core team and approved in the Charter - Breakthrough Next Generation Derivative Design Change |
| - Review Recommended Tailoring | The Tailoring Plan tab of the PDP3 RACI includes tailoring guidelines for each class of project specified in Table 11. The recommended guidelines provide the minimum set of deliverables recommended based on the classification of the project. Some deliverables are “Mandatory” and must always be produced. |
| - Perform Project-Specific Tailoring | Determine the appropriate tailoring of phases, deliverables, and reviews by – Adding or removing entire deliverables. Changing scope of individual non-mandatory deliverables. Adding, removing, or combining reviews. Combining phases. - It is especially important to add deliverables that are specific to the project but that don’t appear in the PDP3 RACI. For example, development builds and prototype iterations are critical to the project but are project-specific and therefore not listed in the PDP3 RACI. Unless these deliverables are added, the project scope and schedule will be less than what is actually required. - A rationale / explanation must be provided for all PDP “Recommended” deliverables that are tailored-out of the project. These rationales are subject to review and approval as part of the project plan review. - All deliverables listed as “Mandatory” cannot be tailored out and must be kept in-scope. - Sometimes, a substantially equivalent work product already exists (e.g., from an earlier product version). These can be leveraged with detailed rationale in the Project Plan or a separate Adoption Justification referenced in the DHF. In these instances, the deliverables are not considered to be “tailored out” but rather replaced by the equivalent work product. |
| - Complete Tailoring Plan | A Tailoring Plan template, included in the PDP3 RACI, is populated to document all tailoring decisions proposed by the Core Team. This template includes all deliverables to be produced in each phase of the project as well as the rationale behind each “Recommended” deliverable tailored-out or reduced in scope. |
| - Review / Approve Recommended Tailoring | The Tailoring Plan is approved as part of the Project History File (PHF). |

### Tailoring Examples

The Tailoring Plan tab in the PDP3 RACI shows a list of deliverables and reviews that are mandatory for all project types and a set of additional recommended deliverables for each project type. This combined set provides a starting point for tailoring.

For Breakthrough and Next Generation projects, less tailoring is expected since PDP3 was designed to accommodate these more complex projects. For Derivative and Design Change projects, more tailoring is expected to ensure only value-added activities are included.

Figure 5 – Example of Section of Completed Tailoring Plan

### Combining Phases

As part of Tailoring, the team may choose to combine PDP Phases.  Rationale and justification must be approved as part of the Project Charter.  It must also be addressed in the Project Plan.

Phase combination standards:

- Though Phases may be combined, all Mandatory deliverables must be completed.  (Whether created or adopted.)

- In almost all cases, Phase 2 and 3 are not combined.  If they are, special allowances to properly document Design Freeze must be addressed in the Project Plan.

- If Phase 4 and 5 are combined, Mandatory deliverables for Phase 5 must be addressed prior to product launch.

- Combined Phase Reviews must cover content from all of the reviews that are combined.

### Tailoring Phase and Design Reviews

PDP3 Phase and Technical Reviews can be combined to fit the needs of the project.  Combined reviews cover content from all combined reviews in one session.

Table 12 provides recommendation guidance on possible combination of reviews.  Phase Gate reviews are combined only when the phases themselves are combined.  A “Stand-alone” recommendation suggests not combining with any other review content.

Table 12 – Recommended Tailoring of Reviews

| Reviews | Breakthrough | Next Gen | Derivative | Design Change |
| --- | --- | --- | --- | --- |
| User Needs Review | Stand-alone | Stand-alone | May combine with subsequent reviews | May combine with subsequent reviews |
| Concept Selection Technical Review | Stand-alone | Stand-alone | May combine with Phase 1 Gate Review | May combine with Phase 1 Gate Review |
| Phase 1 Gate Review (Concept Phase Exit) | Stand-alone | Stand-alone | May combine with Phase 2 Gate Review | May combine with Phase 2 Gate Review |
| Technical Specifications Review | Stand-alone | Stand-alone | May combine with Design Freeze Technical Review | May combine with subsequent reviews |
| DFX Technical Review | Stand-alone | Stand-alone | May combine with Design Freeze Technical Review | May combine with subsequent reviews |
| Design Freeze Technical Review | Stand-alone | Stand-alone | May combine with Phase 2 Gate Review | May combine with Phase 2 Gate Review |
| Phase 2 Gate Review (Feasibility Phase Exit) | Stand-alone | Stand-alone | May be combined with Phase 1 Gate Review | May be combined with Phase 1 Gate Review |
| V&V Technical Review | Stand-alone | Stand-alone | May be combined with Phase 3 Gate Review | May be combined with Phase 3 Gate Review |
| Phase 3 Gate Review (Development Phase Exit) | Stand-alone | Stand-alone | May be combined with Clinical Readiness Review or Phase 4 Gate Review | May be combined with Phase 4 Gate Review |
| PPQ Readiness Technical Review | Stand-alone | Stand-alone | Stand-alone | May be combined with Manuf. Transfer Review (but performed prior to PPQ start) |
| Manufacturing Transfer Technical Review | Stand-alone | Stand-alone | Stand-alone | May be combined with PPQ Readiness or Phase 4 Gate Review |
| Clinical Readiness Review | Stand-alone (only needed if human use in Phase 4) | Stand-alone (only needed if human use in Phase 4) | May be combined with Phase 3 Gate Review (only needed if human use in Phase 4) | May be combined with Phase 3 Gate Review (only needed if human use in Phase 4) |
| Phase 4 Gate Review (Qualification Phase Exit) | Stand-alone | Stand-alone | May be combined with Phase 3 Gate Review. May be combined with Phase 5 Gate Review in rare instances. | May be combined with Phase 3 Gate Review. May be combined with Phase 5 Gate Review. |
| Lessons Learned Technical Review | Recommended review; Stand-alone | Recommended review; Stand-alone | Recommended review; Stand-alone | Recommended review; Stand-alone |
| Manufacturing Transfer Update Review | Recommended, review; May be combined with Phase 5 Gate Review | Recommended, review; May be combined with Phase 5 Gate Review | Recommended, review; May be combined with Phase 5 Gate Review | Recommended, review; May be combined with Phase 5 Gate Review |
| Phase 5 Gate Review (Project Close-Out) | Stand-alone | - Stand-alone | - Stand-alone. May be combined with Phase 4 Gate Review in rare instances. | - May be combined with Phase 4 Gate Review |

### Dealing with Previously Developed Products

When PDP3 will be applied to products not developed using PDP3, the tailoring guidelines can be used to identify gaps in the documentation between the process used to develop the product and PDP3, such that the gaps can be addressed and risks mitigated.

A gap analysis can be performed by reviewing all mandatory deliverables, as defined in the PDP3 RACI, and for each mandatory deliverable –

- Identifying if an equivalent document is available.

- Determining if the information required by the deliverable is contained within any other document.

- Determining if the deliverable is applicable for the type of project planned.

- Using the Tailoring Plan Template, create a cross-reference between PDP3 deliverables and the deliverables for the previous product that satisfy the PDP3 deliverables.

- Address any remaining gaps.

# Appendices

[THIS PAGE INTENTIONALLY LEFT BLANK]

# Appendix A – Process Considerations

This section is designed to provide additional best practices and process guidance for critical elements of the PDP.  These practices are recommended and provide confidence that a quality product is delivered that meets both the customer and business need.

Since each product is unique, not all elements of the Process Considerations section will apply to all projects.  It is important for the Project Manager and team to discuss how to apply these best practices such that the final product meets both the customer and business needs.

This Process Consideration section is for reference purposes only. Please refer to Covidien’s quality system documentation (e.g., QAP4.1) for mandatory design control requirements.

[THIS PAGE INTENTIONALLY LEFT BLANK]

### A1 – Gathering the Voice of the Customer

One of the most important activities performed during the Concept Phase is gathering the Voice of the Customer (VOC).  Collection of good data from our customers early on has been proven to be imperative to creating a commercially successful product.  Figure 6 depicts the key VOC activities –

Figure 6 – VOC Activities

Creating the Voice of the Customer Plan – The VOCP details all activities that will be performed throughout the product development cycle that are used to gather the Voice of the Customer.  This includes all “touch points” where customer feedback on the product is being sought, as well as the anticipated number of iterations through the VOC cycle documented in Figure 6.

Creating the user needs – Translating the user needs from VOC to more specific user needs that define the features and characteristics of the product.

Prioritizing the requirements – Determining which requirements are most important to the customer. This priority will be used by Engineering to negotiate the technical specifications with marketing if the need arises.  The Critical to Quality (CTQ) requirements are the critical few requirements, typically 3-5 total, that are the highest priority.

Reviewing the user needs – The review of the user needs ensures that the requirements represent the user needs and meet the characteristics of good requirements, especially that they are complete and can be validated.

Translating user needs into technical specifications – Translating the User Needs into the Technical Specifications re-specifies the requirements using “engineering language”.  The technical specifications add requirements that don’t originate with the customer (e.g. national or regulatory standards), and would include limits and tolerances where applicable.

Translating technical specifications into engineering design specs – Adequate design documents must be produced to support the development of Concept prototypes. While neither complete nor formally reviewed, these specs should adhere to good engineering practices.

Validating User Needs with the customers – Validation of the requirements with the customer is performed to ensure that the Voice of the Customer is accurately captured.

- Validation of user needs is not the same as product validation, which will occur after the product is built and in the Qualification Phase.

- One typical and effective way of validating requirements with the customer is through the use of various prototypes to display important product characteristics.

Several iterations of the activities above may be performed throughout the Concept Phase to develop preliminary User Needs and Technical Specifications adequate to begin the Feasibility Phase.

### A2 – Lean Product Development

Elements of the Lean Product Development methodology have been integrated into the PDP3 as specified in Table 13.

NOTE – Reference Lean Terminology in Appendix A3 for a definition of terms in the “Going Leaner” column.

Table 13 – Lean Principles in PDP3

| Lean Principle | PDP 2.0 Lean Activities | Going Leaner |
| --- | --- | --- |
| Define What Adds Value to the Customer | VOC activities are the starting point for defining what adds value to the customer. User Needs flow down to the manufacturing processes that are required to build the product, and can help in determining where the non-value added activities in the process are. | Identify Muda (see definition below in A3), and remove if non-essential. |
| Explore Alternative Solutions | During the Concept Phase, many alternatives are explored and the design begins to converge into an optimum solution. | Create trade-off curves and decision matrices to help evaluate the multiple alternatives. |
| Eliminate Waste | PDP3 provides for tailoring the process so that only value added activities are included. | Create Value Stream Maps (VSM) of the present and future state to identify wastes. The Nine Wastes of Product Development section in Appendix A13 identifies the types of wastes typically observed in a product development process. Project teams are encouraged to seek out and eliminate these wastes in their implementation of PDP3. Two methods of eliminating waste (see definitions below in A3) are – Use of an Obeya Poka Yoke |
| Continue to Learn and Improve | The Concept Selection Report details the knowledge gained about the competition. PDP3 includes a Lessons Learned Technical Review to reflect on the performance of the project. | Use of A3 problem-solving methods. Encourage team to question everything by asking why, why, why, why and why to drill down to the real cause of problems. |
| Reduce Variation by Standardizing Processes, Parts, and People | PDP3 includes many activities that are designed to reduce variation in the design and in the process. PDP3 incorporates the re-use of components as a method of evaluating the competing concepts during the Concept Phase. | Reduce variation in the design by re-use of components between products. Reduce variation in the process by re-using processes that are capable and in control. Reduce variation in engineering skills by providing training. |
| Integrate Suppliers Into the Product Development Process | Supplier Development activities occur throughout the product development process. | Value suppliers for their technical expertise. Balance material cost and quality. |

### A3 – Lean Terminology

There are various terms that are used in Lean Product Development, including –

Muda: Non value-added – Any activity that adds time or cost to the product at no benefit to the customer.

Obeya: Big room – In Lean Product Development, each project has its own space where stakeholders can go to be visually updated of project progress.  The Obeya provides visual updates by –

- Displaying charts and graphs depicting program timing, milestones, and progress to date.

- Displaying work products that are indicative of the progress of the project.

- Providing a place for the Project Team to hold daily 5-minute meetings to discuss current issues and activities.

Poka Yoke: Error proofing – Error proofing (prevention) helps eliminate waste from the process by reducing inspection time.  Error proofing can include –

- Design of fixturing/process to prevent non-conforming product and prevent errors from being made and passed on to the next process.

- Design of automatic feedback and control processes to prevent non-conforming product from being made or passed on to the next process.

- Use of checklists.

- Re-using components across products.

- Re-using processes that are capable and in control.

A3 Analysis: Single Sheet of Paper – A3 is a standardized way of communicating by creating a report using one side of a piece of paper (A3 refers to the paper size). There are three types of A3 reports –

- Status – An A3 status report tells the story of an ongoing activity and answers the questions –

- What is the objective of the activity?

- How well are we doing meeting the objective?

- How much was accomplished versus what was planned?

- What issues need to be addressed?

- What do we plan to do next?

- Informational – An A3 informational report is used to tell a story with the intent of sharing information.

- Problem Solving – An A3 problem solving report is used to tell the story of an approach to solving a specific problem.  It should contain the information using a Plan-Do-Check-Act methodology.  The main elements of an A3 problem solving report are –

- Description of the problem.

- Root cause analysis of the problem, using the “5 Whys” to ensure drill-down to the root cause.

- Short term and long term corrections to the problem.

- Plan for implementation of the corrections (what, where, who, when).

- Follow up actions to verify that the corrections were successful.

Where possible, A3 reports should use visual aids to tell the story instead of words.  Every word included in the A3 report should be necessary to help tell the story.

#### Value Stream Maps (VSM)

The Value Stream Maps (VSM) are useful because –

Waste and sources of waste can be visually identified.

Instead of fixing just one issue, the picture of the entire future state is revealed, addressing many issues.

An initial value stream map captures the present state (what we are doing today) and a second VSM captures the future state (where we would like to be). Activities are then chosen to incrementally achieve that future state at some point in the future. The path to the future state becomes an iterative process of continuous improvement.

#### Nine Wastes of Product Development

There are nine wastes that a Lean Product Development process such as PDP 2.0 attempts to eliminate.  The nine wastes are –

- Scatter - organizational changes or events that disrupt teamwork, the flow of information and knowledge, and the cadence or rhythm of a development project (e.g., reorganization, changes in work schedule, etc.).

- Barriers to Communication – Intellectual class structures and departmentally focused thinking (e.g., language, distance, information systems, etc.).

- Wrong Tools – Focusing more on the tool instead of the best way to acquire the knowledge and learning needed.  Some examples are –

- Prototyping versus simulating.

- Simulating versus testing.

- Re-inventing versus re-using.

- Handoff – Separation of knowledge, responsibility and feedback.

- Useless Information – Time spent by product development staff on preparing information, progress reports, studies, analyses and presentations that add no value to the final design of the product.

- Waiting – Conventional product development activities are organized sequentially, which can force decisions and learning to stack-up and happen at once.  For example –

- Waiting for signatures.

- Waiting for technical information.

- Wishful Thinking – Setting specifications early before enough data is available to support the concept or design.

- Testing to Specifications – Testing conformance to specifications with the limited goal of passing the test. Testing should also be performed for other reasons, e.g. discovering how to make the product fail beyond the conditions of use, how to make it better or cheaper, etc.

- Discarded Knowledge – Most knowledge generated during a development project is unintentionally thrown away because development teams are disbanded, documentation is incomplete or inaccessible, etc.

Reference section in Appendix A13 on Eliminating Waste with Lean Manufacturing for additional information on Lean.

[THIS PAGE INTENTIONALLY LEFT BLANK]

### A4 – Design Convergence of Different Concepts

The Concept Phase evaluates various concepts and selects the concept that is the best strategic fit for the customer and the business.  In other words, in this phase, product designers must obtain a good understanding of the problem domain (user needs) and start exploring the boundaries of the solution domain (alternative designs).  The goals of the phase are met when it is proven that one or more solutions exist to solve the problem.

The Concept Phase starts with a set of user needs and develops several ideas on how to best address those needs. These ideas go through a series of quick iterations using concurrent engineering to evaluate the alternatives from several viewpoints (R&D, Manufacturing, Marketing, Quality, etc.).

When a concept is shown to have a low probability of success, continued iterations are halted on that concept.  Areas that show promise can be “converged” with other concepts until design convergence is achieved and one idea is selected to continue to Feasibility, or until no reasonable alternative exists.  The number of iterations can vary for each concept and for different projects.

#### Typical Iteration

Figure 7 depicts a typical flow for the design convergence iterations. After the Evaluate Prototype step, a decision is made whether to –

- Continue to improve the concept.

- Develop another concept.

- Stop iterating, with or without a suitable concept.

Figure 7 – Typical Iterations in Concept Innovation Phase

#### Evaluating the Concepts

In order to properly and efficiently evaluate the various concepts, the following tasks are performed –

Identify the CTQs and other key metrics to help evaluate the concepts.  Key criteria that define success in these areas are identified in Table 14 below.

Iteratively refine the concept design and compare it against the CTQs (i.e. user needs) and key metrics (i.e. business needs).

Eliminate concepts that indicate a high probability of not meeting customer or business needs and continue to refine concepts that have favorable metrics.

Converge and combine successful design elements and re-evaluate against the CTQs and key metrics until one or more concepts are selected.

Determine what technical hurdles are encountered and where risks remain.  These risks are documented and mitigated through the use of the Technology Risk Assessment (TRA) within the Project Plan.

Table 14 – Criteria for Evaluating Competing Concepts

| Key Metric | Evaluation Points |
| --- | --- |
| Intellectual Property | Does the design have patent issues that would need to be resolved? Does the design include technologies that need to be licensed? Is the design likely to result in a new patent or copyright? |
| Manufacturability | Would manufacture of the product require new processes or significant process development time? Can the product be manufactured for the target cost? Would manufacture of the product require domain expertise not currently available to the company? |
| Reliability | Does the design require components for which the company does not have historical data? Does the design require new technology for which there is no historical data? Does the design use existing technology, but in a new, unique, or different environment? |
| Technology | Does the design involve technology new to the company? Are the required technologies within the company’s core competencies? Is outsourcing development or manufacture of the design required? Is there potential for design reuse? |
| Safety | Does the design introduce any potential new hazard? Can all foreseeable hazards be mitigated adequately? |
| Regulatory | Does the design require a 510(k)? Does the design require an IDE submission? Does the design add any additional regulatory risk? |
| Clinical | Would the design require a clinical study? If so, how large a study and what duration? |
| Sourcing | Is there a potential to use existing suppliers? Does the design require the qualification and development of new suppliers? |
| Competitors | Does the design have any characteristics that would differentiate it from the competitor’s product? Can the design compete in the market? |
| User Needs | Does the design meet or exceed user expectations? Are there any user needs that are not met by the design? Would the design be accepted by the customers? |
| Critical to Quality Requirements | Does the design meet all the Critical to Quality (CTQ) requirements? |

[THIS PAGE INTENTIONALLY LEFT BLANK]

### A5 – Characteristics of Good Requirements

Definition – A requirement is a property that a product must have to provide value to a stakeholder. [IEEE]

Good requirements are an essential part of any product development process, and they set the foundation on which subsequent product development activities are built.  The foundation must be strong if the product and project are to succeed.

Requirements documents use different words to distinguish requirements, goals, and other statements –

- Shall – Indicates a requirement.

- Should – Indicates a goal.

- Can, Will or May – Indicates a fact or purpose.

Requirements are stated in an active voice using the word “shall”.  For example:

The catheter shall weigh less than .25 lbs.

Table 15 defines the characteristics of good requirements.

Table 15 – Characteristics of Good Requirements

| Characteristic | Description |
| --- | --- |
| Uniquely Identified | - Each requirement shall have a unique identifier that does not change. |
| Unambiguous | Each requirement should have only one interpretation. Each major characteristic of the final product shall be described using a single unique term. If a term can have multiple meanings, then it shall be defined in a glossary. The following ambiguous words should be avoided when writing requirements – - Usually - Slow / Fast - Frequently - Approximately - Effective - Good - Easy - Always / Never - Normal - Average - Adequate - Possible - Instantaneous - Infrequent - Higher / Lower - Better |
| Correct | Each requirement shall be compared with predecessor documents and standards to ensure that it is in agreement and technically feasible. |
| Verifiable | Each requirement specified shall be testable. A testable requirement is one for which measurable acceptance criteria can be provided to verify the requirement through test, inspection, analysis, or review. |
| Complete | - Each requirement specified shall express a whole idea or statement. |
| Consistent | The requirements specification does not have requirements that contradict any other requirement in the document. Consistent terminology is used throughout the document. |
| Modifiable | Each requirement in the requirements specification shall be expressed separately, not intermixed with other requirements. The same requirement shall not appear in the requirements specification in more than one place. |
| Traceable | The requirement shall be backwards traceable to higher-level requirements, forward traceable to lower level requirements, design, implementation and verification. |
| Prioritized | Each customer requirement shall be ranked for importance based on the value to the customer. Typically, three priority levels are used – - High – The requirement must be implemented in the finished product. - Medium – The requirement should be implemented in the finished product, and reasonable effort should be invested to include it. - Low – The finished product will be successful without this requirement, but the value of the product would be enhanced by including it. |

### A6 – Developing a Robust Design

Definition – A robust product is one that works as intended regardless of variations that are encountered in the manufacturing process, variations resulting from deterioration over time, or variations in use related to people or environment.

In order to achieve a robust design, a Design for Excellence (DFX) methodology is used.   Design for Excellence contributes to a robust design by encouraging the use of concurrent engineering principles throughout the development of the product. Use of DFX results in improved customer satisfaction because of –

- Improved product quality

- Improved reliability

- Reduced product cost

DFX is also interpreted as a collection of “Design for” methodologies.  In PDP3 the “X” stands for:

- Manufacturability and Assembly

- Reliability

- Safety

- Six Sigma

These methodologies are summarized in the sections that follow.

#### Design for Manufacturability and Assembly

Design for Excellence (DFX) techniques are used in the Feasibility Phase to simplify the design of the product and make it easier to assemble and manufacture.

A DFX technical review is performed in the Feasibility Phase to –

Review the current state of manufacturing capabilities.

Identify critical suppliers.

Identify opportunities to use existing suppliers.

Identify opportunities to reduce the number of parts.

Identify opportunities for design simplification.

Identify tolerances that might not be within the capabilities of the manufacturing processes.

Review tolerance stack up of any components that may be reliant on tolerances or capability of assembly equipment

Identify assembly difficulties and address assembly improvements through the design of the components.

Estimation of Costs of Goods Sold (COGS) and assembly costs.

Proposed changes to improve the manufacturability of the product based are documented in the DFX Technical Review.

#### Design for Reliability

Design for Reliability (DFR) techniques are used in the Feasibility Phase to identify sources of variation that can impact the design and to eliminate or reduce the impact of sources such as:

| Source of Variation | Description |
| --- | --- |
| Tolerances | Variation allowed by the device design due to the tolerances and tolerance stack-up. |
| Piece to piece variation | Variation due to the manufacturing process. |
| Deterioration over time | Wear and tear as the part is used over its expected life. |
| Use duty cycle | Variations due to continuous or intermittent use. |
| Operating environment | Variations due to the user and / or environment. |
| Interactions / Interfaces | Variations due to other entities that interact or interface with the device. These entities can be materials, components, people, or other products and systems. |

Figure 8 contains a graphic representation of how PDP3 deliverables are used together with DFR techniques to ensure robustness in the design.

Figure 8 – Robust Design DFR Techniques

As shown in Figure 8, the principle Feasibility Phase activities that contribute to a robust design are described below –

[1] Technology Risk Assessment (TRA) – The TRA is documented within the Project Plan and is an analysis of the potential technical threats to achieving the project objectives.  Although the purpose of this document is to help manage project risk, some of the technical risks identified can be assessed and mitigated by performing robust design activities.  Areas of risk include:

- Failure to meet standards

- Failure to meet performance requirements

- Intellectual property

- Product complexity

- Use of new technologies

- Use of new materials

- Manufacturability issues

- Knowledge gaps

[2] Specification Conformance Matrix – User Needs and the Technical Specifications define the intended uses of the device and specify what functionality the device must exhibit.

[3] Design Characterization Matrix – This matrix defines the activities and testing needed to develop the Design Drawings, as well as the design challenge testing needed to ensure that those specifications are correct.  Design challenge is defined below.

[4] Design Drawings – The Design Drawings document the design that implements the customer and product requirements.

[5] and [10] Manufacturing Process Specification – The [5] Manufacturing Procedures/Quality Procedures document the processes that are used to produce the design.  Once the sources of variation and failure are understood, the [10] Manufacturing Procedures/Quality Procedures are updated with methods of control.

[6] Design Challenge Testing – Design Challenge Testing is performed to understand the sources of variation for the product.  Design Challenge Testing can include testing of –

- Component dimensions

- Component material

- Subassembly dimensions

- Top assembly dimensions

- Functional requirements of device design

- Extreme environments for use

- Development of extreme test models to challenge the device at and beyond its intended design limits and to improve the signal-to-noise ratio in order to amplify failure rates.

Design Challenge Testing is performed as per the Design Challenge Testing Protocol(s) and test results are summarized in the Design Challenge Testing Report(s).

[7] Tolerance Analysis – Tolerances specified on drawings can impact both the design and the manufacturing processes.  Tolerance Analysis also contributes to identifying sources of variation.

Figure 9 – Tolerance Impact

Tolerance Analysis is performed by determining the worst case tolerances and summing the component tolerances linearly. Each component dimension is assumed to be at its maximum or minimum limit, resulting in the worst possible assembly or performance limits.

Tolerance Analysis is performed on dimensional tolerances of all interfaces where two components interface including –

- Components

- Subassemblies

- Top assemblies

- Packaging

- Accessories

Monte Carlo simulations can also be used to perform tolerance analysis.  Monte Carlo analysis is useful in cases where more than three tolerances are stacking or in complex non-linear cases.  Monte Carlo simulation can also be used to estimate mechanical reliability in complex mechanical systems or to optimize complex scenarios to determine the optimum tolerances for reliability or cost.  This technique is described in the Robust Design Techniques section.  The results of the tolerance analyses are documented in the Tolerance Analysis Report.

[8] Risk Management Workbook (RMW) – RMW includes the Procedure Hazard Analysis (PHA), the User Failure Mode Effects & Criticality Analysis (UFMECA) and the Design Failure Mode Effects & Criticality Analysis (DFMECA).  The RMW identifies potential hazards of use and potential failure modes for the product, assesses the risk associated with those hazards or failure modes, prioritizes issues for corrective action, and implements corrective actions to address the most serious concerns.

[9] PFMECA – PFMECA is used to determine how the process can contribute to product failures.  The PFMECA is a section within the Risk Management Workbook (RMW)

[11] Design V&V – Design Verification & Validation is performed to verify that controls established to minimize variability and mitigate failures are adequately performing their intended purpose.

[12] Process Validation – Process validation is performed to establish process controls.

[13] Marketing Support Material and Testing – Documents the Reliability testing performed to verify Marketing claims.

[14] Launch Plan – The Launch Plan defines the Marketing Support Testing that needs to be performed.

#### Design for Safety

A design cannot be robust without first ensuring safety.  Risk Management activities are integrated throughout the phases of PDP3, but these activities are of particular importance in the Feasibility Phase where the design can still be influenced to ensure an effective and safe design.

Figure 10 shows how risk control measures are integrated into the PDP process steps that ensure hazards are assessed and controlled.  Knowledge obtained from the safe design elements are fed back into the design.  This process iterates until the design is found to have an acceptable level of risk.

The safe design elements included in PDP are intended to provide –

An inherently safe product.

Protective measures for known and foreseeable hazards that have not been eliminated through design work.

Information to further mitigate remaining hazards.

Figure 10 – Safe Design Elements

#### Design for Six Sigma (DFSS)

Six Sigma is a methodology for problem solving and continuous improvement.  DFSS is used to improve products and processes by identifying and understanding the potential sources of variation and then making the design and the process less susceptible to these sources of variations as long as the costs incurred do not exceed the perceived benefit.

In order to identify and control variation, DFSS uses the Define, Measure, Analyze, Improve, Control (DMAIC) problem solving method to affect improvement.  DMAIC means –

- Define – Describe the problem in a problem statement.

- Measure – Determine the current performance or capability of the process.  Design of Experiments (DOE) are a very useful tool to measure current performance and establish causes of variation.

- Analyze – Perform statistical analysis of the data to understand the cause and effect relationships in the design or the process.

- Improve – Implement and verify improvements based on the analysis of the data.  DOE can also be used to determine that the causes of variation have been reduced.

- Control – Plan to ensure improvements are sustained.

There are various tools that can be used to incorporate DMAIC.  Table 16 defines tools that can be used for improving the product and the process.

Table 16 – DFSS Tools

| DMAIC | Product | Process |
| --- | --- | --- |
| Define | - Problem Statement | - Problem Statement |
| Measure | - DOE | - Process Maps - Gauge R&R |
| Analyze | - Affinity Diagrams - Fishbone Diagrams - RMW - Statistical Analysis - Analysis of Variance (ANOVA) - Tolerance Analysis - Regression Analysis - Reliability - Descriptive Statistics | - Capability Studies - PFMECA |
| Improve | - DOE | - DOE |
| Control | - N/A | - Statistical Process Control (SPC) - Process Control Plans - Poka-Yoke (mistake proofing) |

#### Robust Design Techniques

There are many techniques which can be used to ensure that the design is a robust design.  Two well established techniques that are recommended are –

Monte Carlo Simulation – Monte Carlo simulation is used to perform tolerance analysis.  Monte Carlo simulation takes the individual distribution of the tolerances of each individual component involved in a tolerance stack-up.  The simulator randomly selects values within these distributions and incorporates them into a mathematical model representing the tolerance stack up situation.  The mathematical model is generated by the engineer based on the specifics of the tolerance situation.  The output is only as good as the information input and model generated by the engineer.  The simulator then generates an output distribution.  The results of the simulation include the population mean and standard deviation and normality, the 3-sigma limits for the worst case tolerances. This analysis provides the percentage of components estimated to be in and out of the specification range (estimated design or process capability).

Design of Experiments (DOE) – DOE is used to identify important variables, to estimate their effect on a certain product characteristics, and to optimize the settings of these variables to improve the design robustness.  The key steps in designing an experiment are –

- Identify the objective of the experiment. Typically these include finding true causes of problems, finding how causes interact, finding the best solution to a problem, or testing a solution to ensure it has no undesirable side-effects.

- Define what is to be measured to show the results of the experiment.

- Identify the factors that are to be controlled during the experiment. Consider all things that may affect the measured result, then select those that are to be varied and those that are to be held constant or otherwise monitored.

- For each factor selected, identify the set of levels that the experiment must consider.

- Select the actual trials to take place.

- Perform the trials as planned.

- Analyze the results.

- Take appropriate actions based on the results.

[THIS PAGE INTENTIONALLY LEFT BLANK]

### A7 – Developing a Robust Process

A robust process is one that can produce product that consistently meets specifications within target cost and yields.  In order to achieve a robust process, the potential sources of variation must be understood and controlled.

Figure 11 displays the sources of variations that can affect a process, and therefore impact the performance of the product.

Figure 11 – Process Variation

The sources of variation that affect a process include –

Variation due to materials.

Variation due to equipment (gauges).

Variations due to the environment.

Variations due to operator.

Variations due to the methods used.

PDP3 has activities designed to determine –

How the different sources of variation can influence the output of the process.

How interaction of the various inputs can affect the output of the process.

How to monitor and control the impact of variation.

How the process is performing against established criteria.

Statistical analysis is extensively used to identify and control process variables.

#### Process Characterization & Optimization

Several methods are used to characterize and optimize the process as it evolves in the Development Phase –

Design of Experiments (DOE) is used to –

- Determine the interaction between variables. Determine which variables are critical and not critical for control of the process.

- Define the optimum input settings and the operating window for the process to produce components and/or assemblies within specifications.

- Test a solution to ensure that there are no undesirable effects.

- Determine the causes of variation.

Process flow diagrams are created to –

- Define and analyze processes.

- Build a step-by-step picture of the process for analysis, discussion, or communication.

- Define, standardize or find areas for improvement in a process.

Fishbone diagrams are used to determine key sources that contribute most significantly to the problem being examined. DOEs can then be used to verify and quantify the contribution and determine possible corrections.

Process Failure Mode Effects and Criticality Analysis (PFMECA) is performed to ensure that potential failure modes and their associated manufacturing or assembly process causes have been considered and reduced as much as is technically and economically feasible.  The PFMECA also –

- Facilitates prioritizing preventive and corrective actions.

- Identifies variables that can be subjected to process controls.

#### Process Capability

Once process characterization has been completed and the process has been optimized, the process capability of each process is determined.  A process capability index relates the engineering specification to the observed process behavior.

When calculating a process capability index it should be noted that a sufficiently large sample size must be used, and an assumption of a normal distribution is made.

Cpk=min⁡[USL- μ3σ, μ-LSL3σ]

Where:  Min = minimum; USL = Upper Specification Limit;

LSL = Lower Specification Limit; μ = Mean; σ = Standard Deviation

Typically Cpk values of 1.33 or higher are considered acceptable and indicate a stable process.

| CpK | Sigma Level |
| --- | --- |
| 1.0 | 3 |
| 1.33 | 4 |
| 1.66 | 5 |
| 2.0 | 6 |

Control Charts are used to determine if the process is in a state of statistical control and whether the process has the capability required to produce product within specification.  Control Charts –

Provide visibility when special causes of variation are present.

Reflect the extent of common causes of variation that must be reduced.

#### Gauge Repeatability & Reproducibility

Before a manufacturing process can be controlled or optimized, the uncertainty of the measurements being made as part of that process must be understood.  A Gauge Repeatability & Reproducibility (R&R) study is performed to determine the precision of measurements using instruments and gauges.  The precision of the measurement has two components –

Repeatability – The precision of the measurement under similar conditions, i.e., the difference in the average of the measurements made by the same operator using the same gauge measuring the same part multiple times.

Reproducibility – Precision under different conditions, i.e., the difference in the average of the measurements made by different operators using the same gauge and measuring the same part.

Figure 12 – Repeatability vs. Reproducibility

A typical procedure for conducting a Gauge R&R study includes the following steps –

- Obtain samples that represent the actual or expected range of process variation.  Typically, samples from three different lots are used, and the total number of samples should be more than five.  Each part needs to be identified for the study.

- Identify the number of operators to be used in the study.

- Setup the equipment as required by manufacturing procedures, including calibration.

- Each operator measures all parts in random order.

- Repeat the cycle for the number of trials required.

- Analyze the data to determine –

- If the measurement system has sufficient resolution to detect part-to-part variation.

- If the process is in control.

- If the measurement process can be applied consistently by each operator.

- How much of the total variation can be attributed to the gauge, operator, and parts.

### A8 – Verifying and Validating the Product Design

As Figure 13 graphically depicts, Design Validation ensures that the product actually does what the customer (or user) wants the product to do in the intended use environment (or a close simulation).  Design Verification ensures that the product meets its design requirements.

Figure 13 – Design Verification and Design Validation

#### Verifying the Product Design

Definition – Design verification means confirmation by examination and provision of objective evidence that specified requirements have been fulfilled.  (FDA)

Verification activities are undertaken to provide objective evidence that design outputs conform to design inputs.  The objectives of design verification include –

- Finding defects.

- Assessing conformance to specifications.

- Proving conformance to regulations.

- Verifying safety mitigations are effective.

In order to accomplish these objectives, several key elements of design verification should be incorporated –

- Verification activities that can be used to provide objective evidence that design outputs conform to design inputs.

- Verification tests that can be used to show conformance to specifications and standards.

- Methods for ensuring that design verification tests are designed to find defects, and not just to show that the requirements are met.

- Ensuring the validity of the verification tests through –

- Appropriate selection of sample sizes.

- Test methods used that can provide valid results.

The following sections discuss these key elements in more detail.

#### Verification Activities

As described above, verification activities are performed to gather objective evidence that design outputs conform to design inputs.   Verification activities include –

- Inspection – Visually noting and recording a specific attribute of the product to confirm compliance with requirements. Inspection is used to verify properties best determined by examination and observation (e.g., color).

- Test – Verifying the performance of a product by measuring its characteristics or attributes, e.g. their extent, size, capacity, amount, or quantity.  The following are typical measurements for hardware devices –

- Demonstration – Providing a known set of input conditions and verifying that the device behaves as expected.

- Analysis – Analysis can include –

- Design review of drawings.

- Review of Historical data and / or documentation to be able to determine that the product provides the specified characteristic or feature.

- Simulation of extreme operating conditions when it is not realistic, feasible, or cost effective to establish those conditions for physical testing.

- Certification – Written assurance that the product conforms to a standard.  Certification is typically performed by a third party with subject matter expertise.

Every test specified for verification shall produce some form of objective evidence as described in the process steps above.

#### Verification Tests

Table 17 lists tests that are typically included as part of design verification along with key considerations that make the test effective and efficient.

Table 17 – Verification Tests

| Test | Description | Key Considerations |
| --- | --- | --- |
| Functional Testing | - Testing based on an analysis of the functional specification of a component or system. | - Testing at extremes of the specification. - Compatibility with – - Magnetic Resonance Imaging (MRI) - X-ray |
| Accelerated Aging | - Testing to demonstrate that the product, in combination with its packaging components, performs efficiently, safely, and effectively throughout its intended shelf life, as indicated by the expiration date on the packaging. - Testing that subjects samples to elevated temperatures for specific periods of time to simulate the effects of real-time aging. - The product can be released to market based upon successful accelerated aging of the package and product that simulates the period claimed for product expiration. | - Always understand the materials in the test product to make sure that they can withstand the chosen test temperature. 55 degrees Celsius is the usual temperature for testing. |
| Real-Time Aging | - Testing performed to substantiate the data generated during the accelerated aging process. | N/A |
| Biocompatibility | - Testing to prove that materials that come into contact with tissue, blood, or skin will not cause any adverse local or system effects. | - Applicable standards – - ISO 10993 – Biological evaluation of devices. - FDA Blue Book Guidance Memorandum #G95-1, includes optional tests to consider. |
| Chemical Testing | - Testing to prove that materials that come into contact with tissue, blood, or skin will not cause any adverse local or system effects. | - Required by China FDA - GB/7 14233.1 – Test methods for infusion, transfusion, injection equipment for medical use – Part 1: Chemical analysis methods |
| Sterility | - The purpose of sterilization is to inactivate the microbiological contaminants on a device and transform the non-sterile items into sterile ones. - The sterilization process is validated to show how effective it is at reducing the level of bioburden below that which is specified. - Testing must also verify that there are no negative impacts on the product or packaging as a result of sterilization. | - Sterility Assurance Level (SAL) must be selected (typically 10-6). - Test for effects of re-sterilization on product and packaging. |
| Reliability Testing | - A test that verifies that he product is properly designed and assembled by subjecting it to stress conditions that accelerate potential failure mechanisms. - A test that shows that required reliability can be attained by establishing a reliability prediction based on verification data obtained. | - Tolerance stack-up. - Reliability demonstration tests. - Accelerated life test. |
| Worst Case Analysis | - A test that emphasizes the most unfavorable operating conditions for the device. | - Consider opposites – - Lowest / highest - Minimum / maximum - Don’t limit the analysis to just two variables. |
| Packaging / Shipping / Storage Tests | - A test that emphasizes extremes of transportation and storage conditions. | - Shock and vibration testing to simulate transport by rail, road, sea and air. - High and low environmental temperatures. - High and low humidity. - High and low atmospheric pressure. - Ingress protection. - Salt corrosion. - Packaging tests – - Seal integrity. - Label integrity. - Dye penetration. - Pouch burst. - Ship testing. |
| Environmental Tests | - A test that determines how well products withstand the stresses likely to be experienced in the typical environment in which the device is to be used. | - Shock. - Vibration. - Thermal cycling. |

#### Robustness Testing

- NOTE – Discrepancies and errors in device function and performance are much more costly to correct later in the development cycle.  Typically, the cost to find defects increases exponentially as the development cycle progresses.

- Therefore, verification related activities should be started early in the Development Phase to identify discrepancies as soon as possible.

Various methodologies can be applied in creating verification tests to ensure that the product design is both robust and efficient.  Some of these methods include –

- Design of Experiment (DOE) – In the Development Phase, DOE is used to identify sources of variability and their effects on certain product characteristics or to reduce the sources of variability.  In the Development Phase, DOEs can be used to demonstrate that the sources of variability do not affect a specific product characteristic.  For example, through the use of DOE, it can be demonstrated that no interaction between the sources of variability will impact the ability of the friction characteristic meeting specification.

- Testing at the Specification Limits – In order to ensure that the product performs as specified in the user environment, testing should be performed at all specification limits.  As Figure 14 graphically represents, there are various performance limits specified for the product –

- Specification limits – Specification limits are established based on actual performance taking into consideration manufacturing process variation.

- Design limits – The Technical Specifications specify design limits which the product must meet.  These are the limits used for testing during verification.

- Destruct limits – Destruct limits are arrived at during design characterization if test-to-failure activities are performed.

Figure 14 – Specification Limits

One or more verification tests must be performed to ensure proper operation at the specification limits.  Possible methods of verifying proper operation at the specification limits are –

- Build product at the specification limits and verify proper performance of the product.

- Build product outside the specification limit, in the destruct limit region, and test for proper performance of the product.

- Perform statistical analysis of the product tested and make population estimates based on mean and standard deviation.

- Perform analysis using historical data, design reviews of drawings, or other type of review to arrive at an engineering judgment that the product will perform properly at the specification limits.

#### Ensuring Validity of Test Results

To ensure that design verification (and design validation) testing provides valid data, several key considerations must be taken into account when designing the tests –

##### Sample Sizes

Sample sizes used for design verification testing must be statistically significant.  Many factors weigh into the decision of how many samples to require –

- Type of test (e.g., quantitative variables or qualitative attributes).

- The sample population.

- The number of samples available for testing.

- The risk factors (e.g., confidence intervals).

##### Test Method Validation

Test Method Validation is conducted to confirm that the methods used for verification testing are fit for the intended use and can provide valid data.  Test methods are validated to quantify each attribute below –

- Accuracy – Difference between the true value and the accepted reference value.

- Precision – Closeness of repeated readings on each other.

- Linearity – Change in bias over the normal operating range.

- Gauge R&R – The combined estimate of measurement repeatability and reproducibility.  Reference the section in Appendix A7 on Gauge Repeatability & Reproducibility for further information on Gauge R&R.

- Range – The measurement range for which the test method is validated.

- Uncertainty – An estimated range of values surrounding the measured value in which the true value is believed to be contained.  The following can contribute to measurement uncertainty –

- Bias – The difference between the observed average of measurements and its true value.

- Calibration – A measurement process that assigns values to the property of an artifact or to the response of an instrument relative to reference standards.  The purpose of calibration is to eliminate or reduce bias in the user's measurement system relative to the reference base.

- Environmental Condition – The effect of ambient temperature, humidity, and other environmental factors on the ability of the test method to measure accurately.

#### Validating the Product Design

Definition – Design validation means establishing by objective evidence that device (product) specifications conform to user needs and intended use(s). (FDA)

Validation is the aggregation of all efforts employed to assure that the design will conform to user needs and intended use given projected variations in –

- Components.

- Materials.

- Manufacturing Process.

- Use Environment.

Validation tests are designed to address the needs of all users and all intended uses.  Validation testing can include –

- Historical comparisons to older devices, e.g. a 510(k) historical database search.

- Scientific literature review.

- Simulated use testing (bench testing).

- Clinical testing (clinical studies, if performed).

- Usability studies.

- Animal studies.

- User preference testing.

#### Validation Tests

Table 18 defines key considerations for each of the validation tests defined above.

Table 18 – Validation Tests

| Validation Test | Key Considerations |
| --- | --- |
| Historical Comparisons to Similar Devices | - Review of documentation available to show comparisons with similar devices that can be used as objective evidence of a working solution to an issue. |
| Scientific Literature Review | - A review of publications, such as trade journals, that report original empirical and theoretical work in the natural and social sciences, and within a specific scientific field to highlight evidence that provides objective evidence of a working solution to an issue. |
| Usability Studies (Human Factors Study) | - Usability studies are a method of testing how the user interacts with the product and evaluates – - Performance – how the user accomplishes the intended use of the product. - Accuracy – what types of mistakes the user makes and how they are made. - Ease of Use – how easy is it for the user to learn to use the product without labeling aids. - Satisfaction – how satisfied the user is with the product. - Usability studies can be performed in a simulated use environment. |
| Animal Studies | - Animal studies are a method of simulating the actual use environment as closely as possible without using human subjects. Animal studies can be used to – - Verify efficacy related requirements. - Human factors evaluation. - Usability studies. |
| Clinical Testing | - Clinical testing is performed, if required by the Clinical Development Plan, to validate the use of the product in the actual use environment and using human patients. - Validation in a simulated use environment may be performed if appropriate regulatory clearances are obtained prior to testing. |
| Simulated Use Testing | - Simulated use testing is performed on the bench using models that simulate actual patients (e.g. medical phantoms). - This type of testing is used to verify user acceptance of the product, work flows, and ease of use. |
| User Preference Testing | - User preference studies are performed prior to completion of the design and are used to evaluate how users react to particular features of the product. |

### A9– Design Transfer

The purpose of design transfer is to ensure that the device design is correctly translated into manufacturing production specifications.  Design transfer includes –

Transferring of the Device Master Records (DMR) to manufacturing.

Approval by Operations of all documents that make up the production specifications after assessing to ensure completeness and adequacy through process validation.

Validating all software used in the manufacturing process or for maintaining quality records for the device.

Design transfer is usually not an activity that is performed entirely at one time, but rather it is performed at intervals throughout the Development and Qualification Phases.

Effective design transfer requires that –

- Design documentation needed to manufacture the device is complete and correct.

- Suppliers are approved and performing as expected.

- Manufacturing processes are well understood.

[THIS PAGE INTENTIONALLY LEFT BLANK]

### A10 –Process Validation

Definition – Process validation is establishing documented evidence which provides a high degree of assurance that a specific process will consistently produce a product meeting its pre-determined specifications and quality characteristics. (FDA)

In order to ensure that a process meets this definition, the following activities are included as part of process validation –

- Process Validation Planning.

- Equipment Qualification.

- Process Characterization and Optimization.

- Operational Qualification.

- Performance Qualification.

- Process Monitoring and Control.

Figure 16 shows how these process validation activities are inter-related and

Table 19 describes each process validation activity.

Table 19 – Process Validation Activities

| Activity | Description |
| --- | --- |
| - Validation Planning | The Process Master Validation Matrix (PMVM) is used to manage the validation activities for the processes that will be subjected to validation. The PMVM includes – - A list of all processes to be validated. - A list of equipment qualifications to be performed. - Traceability to validation protocol for each process. - Traceability to validation report for each process. - Requirements for revalidation of processes. The PMVM is created in the Feasibility Phase and is updated as required in subsequent phases. |
| - Equipment/Process Qualification | Equipment/Process Qualification is performed to ensure that the equipment, as installed, is capable of operating within established limits and tolerances, as well as throughout expected operating ranges as established by the Equipment Requirement Specification for that piece of equipment. Equipment/Process Qualification includes Installation Qualification (IQ), which has been previously defined in the Qualification Phase. |
| - Process Characterization and Optimization | Process characterization and optimization includes activities to – - Identify sources of variability – Identify the expected variability of the processes and allow the specification limits to be set. The specification limits are used to establish acceptance criteria for the Operational Qualification and Process Performance Qualification. - Identify which outputs to measure – Define the relationship between process parameters and measured outputs, including which parameters most directly influence outputs. This information is used to determine what outputs to monitor during the process performance qualification. - Identify which parameters to track – Identify key parameters which are allowed to vary in the manufacturing process. These parameters can be used to challenge the process during the process performance qualification. Reference the Developing a Robust Process section in Appendix A7 for further information on these activities. |
| - Operational Qualification | Characteristics of Operational Qualification (OQ) include – Show that at the limits of the proposed process window in terms of maximum operator variation, maximum raw material variation, and at the process parameter limits, the process is capable of manufacturing components or assemblies that meet the product specifications. Typically executed at the high limits and low limits of the process. By high and low limits, the different process parameters are set to the levels that create the (greatest) amount of variation in the product. For example, setting the parameters on a stretching machine to ensure the shortest and longest output material. To ensure a truly robust process if the operational qualification passes statistical tolerance limits or has a capability of 1.33 or greater indicates that the processes is highly stable and unlikely to produce non-conforming material or assemblies. |
| - Performance Qualification | Performance Qualification (PQ) is divided into – - Process Performance Qualification (PQ) – The purpose of process performance qualification is to demonstrate that the process is reproducible and effective by testing the process to determine whether it is capable of consistently producing a product which meets its specification targeted at specific attributes that must be tested independently of the PPQ due to testing complexity issues. - IQ/OQ should be successfully completed prior to the start of PQ. - Full assemblies, subassemblies or abbreviated assemblies (aka “shorties”) are produced independently of the full PPQ to test the specific product attribute(s). - Challenges must be repeated with enough samples to ensure results are statistically valid. - Product Performance Qualification (PPQ) – The purpose of the product performance qualification is to demonstrate that the process has not adversely affected the product and that the product can meet its specifications. The Product Performance Qualification will vary substantially, depending on intended use and on the technology incorporated into the device. Production volume and batch sizes can impact the statistical techniques used. - A minimum of 3 lots is used to establish consistency and process capability across multiple lots of product. Each lot must be large enough to pass statistical requirements independent of the other lots of product manufactured as part of the PPQ. - NOTE – Product Performance Qualification (PPQ) can be combined with Design Validation. |
| - Process Monitoring and Control | Once a process has been shown to be capable and in a state of statistical control through process validation, it must be monitored and controlled to ensure that it remains in a state of statistical control. The monitoring step provides the catalyst for improving the process, as indicated by Figure 15. Figure 15 – Process Improvement Cycle - Monitoring the Process – As a result of the Process Validation, Upper and Lower Control Limits for each process are established. The performance of the process against these control limits should be monitored on a periodic basis to find out of control conditions. Once an out of control condition has been encountered, work to find and correct any assignable causes can be performed to bring the process back under control. - If only common causes of variation are present, the output of a process has a distribution that is stable and predictable over time. - If assignable causes of variation are present, the process output is not stable over time. - Finding and Correcting Assignable Causes – When an out-of-control condition has been identified, the operation of the process should be studied to determine the cause. The cause should then be corrected and prevented from occurring again, as much as possible. Assignable cause variation can include – - Operator variability. - Defective materials. - Machine malfunction. - Operator error. - Improve the Process – Improve the process by reducing the common cause variation. Common cause variation can include – - Inappropriate procedures. - Poor design. - Poor maintenance of machines. - Environmental contaminants. - Variability of materials. - Human factors. - Variations in electrical supply. - Vibration. - Ambient temperature and humidity. - Normal wear and tear. - Variability in settings. - Poor training. - Analyze the Revised Process – When actions have been taken to eliminate an assignable cause, the effect of those actions should be apparent on the control chart. |

Figure 16 – Process Validation Activities

[THIS PAGE INTENTIONALLY LEFT BLANK]

### A11 – Physician Preference Testing (PPT)

Physician Preference Tests (PPT) constitute the first time the device is used on humans, unless a Clinical Study is performed. PPT, if required to validate marketing claims or fulfill regulatory requirements, is an element that may be conducted as an extended element of Design Validation in phase 4. PPT is not performed to verify safety or efficacy and is not considered part of a clinical study.

PPT is performed to confirm the results of design validation in an actual use scenario.  A PPT Study can evaluate how the customer interacts with the product and can confirm the attributes quantified during design validation and listed below.   If desired, performance of CTQ requirements could be confirmed through a PPT Study, i.e.

- Performance – How the customer accomplishes the intended use of the product.

- Accuracy – What types of mistakes the customer makes.

- Ease of Use – How easy is it for the customer to learn to use the product without labeling aids.

- Satisfaction – How satisfied the customer is with the product.

The following are key elements to consider when designing the PPT study –

- Sample size – Statistically valid sample sizes should be used for testing.

- Physician demographics – Select physician demographics based on the population of physicians that are likely to use the device, and the demographics that are most likely to affect the outcome of the therapy and customer satisfaction with the product.

- Level of experience

- Familiarity with the therapeutic procedure

- Information for Use (IFU) – Where possible, incorporate IFU into the study to evaluate the customer’s understanding of the Indications for Use, Contraindications, Intended Uses, Warning and Cautions.

- Intended Uses – Incorporate all intended uses of the product into the study, or perform a study per intended use.

- Support – Availability of Covidien employees at the case is often key to ensure proper execution and data collection

- Acceptance Criteria – Acceptance Criteria specified in the  PPT Protocol should be based on –

- The device performing as intended, based on the feedback from physicians participating in the study.

- Physician preference scores for some or all CTQ features.

[THIS PAGE INTENTIONALLY LEFT BLANK]

### A12 – Lifecycle Management

Figure 17 displays a typical product lifecycle, and boxes shaded in gray are covered in detail in the various sections of PDP3. However, all elements of the lifecycle must be considered when developing the product design and labeling.

Figure 17 – Product Lifecycle

Key considerations for lifecycle phases not discussed previously include –

Distribution

- Ensure distribution only to trained field personnel and trained physicians.

- Ensure that products can be located anywhere in the distribution channel in the event of a recall.

Customer

- The time period for retaining records of the design and manufacture of the product must meet regulatory and legal requirements and be consistent with the expected life of the product.

- The method for communicating to physicians and patients any potential health issues associated with the product must be available.

Disposal

- Labeling must include any special disposal precautions required by standard or by law.

- Environmental hazards associated with the disposal of the product must be taken into consideration in the labeling.

- Hazards to patients, users, bystanders, or animals related to the disposal of the product must be accounted for in hazard analysis.

- Review of applicable standards that govern the disposal of the product must be an input into requirements documents.

### A13 – Manufacturing Transfer

During development of the manufacturing processes, several activities help ensure that the resulting processes are capable of producing a product that -

Meets cost targets.

Meets yield targets.

Can be produced in the quantities required.

These activities include –

- Employing the methodologies discussed in the Developing a Robust Process section in Appendix A7.

Early collaboration between R&D, Process Development, and Operations in assessing the design for manufacturability. If the design is not manufacturable, then it can increase scrap and rework once it gets to production.

Specific areas that should be assessed prior to the transfer of responsibility to Operations, as defined in the Manufacturing Transfer Plan, include the items listed below –

Cost of Goods Sold (COGS)

Cost to manufacture

Manufacturing cycle time

Scrap costs

Rework costs

Yields

Process capability

PFMECA recommendations

Operator training

Supplier performance

Availability of materials

The Manufacturing Transfer Review compares metrics gathered during the Qualification and Launch Phases to ensure that the requirements set forth in the MTP have been met. There are two cases -

- All Requirements Met – If the metrics are favorable during the Manufacturing Transfer Review, then Operations accepts the transfer of responsibility for manufacturing the product.

- Some Requirements Not Met – Conversely, if some of the requirements in the MTP have not been met, then Operations may reject the transfer of responsibility until those requirements are met.  Figure 18 graphically displays possible causes for not meeting MTP requirements.

In this case, Process Development will employ the methodologies specified in the Developing a Robust Process section (Appendix A7) until all requirements specified in the MTP have been met.  Additional methodologies that can be used as an aid to improve the manufacturing transfer metrics are –

- DMAIC Techniques – This Six Sigma methodology allows the team to Define, Measure, Analyze, Improve and Control changes to improve the process.  Reference section Design for Six Sigma (DFSS) (Appendix A6) for more information of the DMAIC process.

- Lean Manufacturing Techniques– Eliminating waste by using Lean Manufacturing.  Reference the section on Eliminating Waste with Lean Manufacturing in this appendix.

Both DMAIC and Lean techniques can be used to address common problems as discussed below.

Figure 18 – Manufacturing Transfer Metrics

As Figure 18 indicates, cost reduction, improved yields, and improved throughput can be achieved by –

#### Reducing Costs

Reducing scrap and rework costs –

- Apply Lean Manufacturing and Six Sigma principles to reduce non-value added activities and eliminate waste from the manufacturing process and reduce scrap costs.

- Apply the Six Sigma DMAIC process to help identify the root causes of rework (typically process, equipment or operator) and improve on the causes to reduce rework.

Reducing COGS –

- Eliminate waste, as defined in the Eliminating Waste with Lean Manufacturing.

- Simplify the design by reducing the number of parts or processes required to manufacture the product.

- Improve the relationship with suppliers to reduce purchase costs.

#### Improving Yields

Improving Process Capability –

- Optimize the process limits.  Limits that are too tight can create false failures, and limits that are too loose can create failures downstream.

- Improve or replace equipment and fixturing to increase process capability.

- Monitor the output of the process using Statistical Process Control (SPC) to give an indication of problems before they occur.  This will allow detection, correction, and prevention of process variation that cause defects.

Error Proofing the Process – Error Proofing is a technique for implementing fail-safe mechanisms to prevent a process from producing defects. Use PFMECA to identify possible errors that can be generated by the process and implement approaches to –

- Prevent an error from happening in the first place.

- Provide a warning before the error occurs.

- Stop the operation so the error cannot occur.

- Provide a method for auto-correcting the error.

#### Increasing Throughput

Improving Yields – Refer to the Improving Yields section above.

Reducing Time to Manufacture  – Time to manufacture can be reduced by –

- Operator training.

- Simplification of assembly of the product.

- Implementation of Lean Manufacturing to eliminate the waste of delay and waiting.

Increasing Capacity – Capacity is the maximum amount of product that a manufacturing line can produce per unit of time.  To increase capacity –

- Ensure that there are no bottlenecks in the production process.

- Ensure batch size is optimized for best capacity.

- Add additional equipment or a second production line, if needed.

- Reduce cycle time.

#### Eliminating Waste with Lean Manufacturing

The seven types of waste that Lean Manufacturing attempts to eliminate are –

- Overproduction – Creating more product than is needed.  Overproduction is often used to make up for poor yields.

- Waiting – Waste of an operator’s time when they are waiting for a part or materials to perform a task or make a decision.

- Conveyance (Transportation) – Movement of products, materials, or information.  Manufacturing floor layouts can cause excess transportation.

- Motion – Unnecessary movement of an operator (walking, lifting, reaching, or stretching).

- Inventory – Finished products, parts, and supplies kept in inventory add cost, and they can become scrap if they exceed shelf life.

- Over-processing – Occurs when a step in the process does not add value to the customer.

- Correction (Defects) – Includes inspections to detect defective product and the paperwork caused by quality issues related to scrap and rework.

See also section Lean Product Development in Appendix A2 for more information on Lean methodologies.

# Appendix B – Changes Before and After Launch

Figure 19 displays the process for making changes throughout the product lifecycle.  This process before product launch is different than the process after product launch.

Figure 19 – Design Change Process

## Design Changes Prior to Product Launch

At each phase gate review, the project team’s request to continue into the next phase will be denied if the objectives of the current phase have not been met.

Additionally, if information is learned that would require activities in the previous phase to be performed, the phase of the project can be “demoted” to accurately indicate the project’s progress. Once demoted, the project proceeds forward according to the PDP guidelines for that phase, including completion of technical and phase gate reviews. Examples of this include –

During Feasibility, a concept previously thought feasible is proven infeasible and additional Concept activities need to be performed.

A failure during design verification or design validation can require a design change that moves the project back into the Feasibility Phase.

## Design Changes After Product Launch

Once the product has been distributed for customer use, design changes may be requested via –

Product Surveillance – Monitoring of the product’s performance in the field will provide design change requests related to –

- Improved safety.

- Improved reliability.

- New intended uses.

Corrective Actions and Preventative Actions (CAPA) – The CAPA system will provide design change requests related to –

- Reliability issues.

- Manufacturability issues.

Voice of the Customer (VOC) – The VOC will provide change requests suggested by the customer, including –

- New features.

- New intended uses.

- Customer dissatisfaction.

Following product launch, the implementation of design changes will require that a new project be initiated.  The tailoring guidelines provided in the PDP Tailoring section should be followed to determine how PDP3 should be applied for the new project.

Table 20 identifies the sources of potential change requests and the project types that may result.

Table 20 – Design Change / Tailoring Cross Reference

| Design Change Request Source | Project Type for Tailoring |
| --- | --- |
| Product Surveillance | - Design Change |
| CAPA | - Derivative Product Development - Design Change |
| VOC | - Next Generation Product Development - Derivative Product Development |

# Appendix C – Configuration Management and Document Revision Levels

Definition – Configuration Management (CM) – A discipline applying technical and administrative direction and surveillance to: identify and document the functional and physical characteristics of a configuration item, control changes to those characteristics, and record and report change processing and implementation status, and verify compliance with specified requirements. (IEEE)

Configuration management provides a method for –

- Uniquely identifying a work product.

- Managing changes to the work product.

- Protecting the configuration of a work product.

A development document is configuration managed with increasing rigor as it progresses through development. More specifically –

- Draft Work Products – Preliminary or draft versions of documents are managed with less rigor.  The following requirements apply to these documents.

- Assigned unique document identifier.

- Assigned a version number to track modifications to the draft document.

- Signed by at least the author (see PDP3 RACI).

Approval Prior to Design Freeze – Prior to Design Freeze, approved documents are managed with increased rigor. The following requirements apply to these documents.

- Assigned unique document identifier (i.e. the same identifier as assigned to the draft).

- Assigned a version number to track modifications to the draft document.

- Signed by author.

- Approved by the team per the PDP3 RACI.

After Design Freeze – After Design Freeze, documents are managed with a high level of rigor.  The following requirements apply to these documents.

- Assigned unique document identifier (i.e. the same identifier as assigned to the approved version).

- Assigned a version number to track modifications to the draft document.

- Signed by author.

- Approved by the team per the PDP3 RACI.

- Changes are made only after impact analysis / design change assessment.

- PDP3 uses three levels of document revision control for configuration management.  Table 20 below describes the differences:

Table 21 – Document Revision Levels

|  | “X” Level Release | Numeric Release | Alpha Release |
| --- | --- | --- | --- |
| Example Revision Numbers Format: | DWGS50467 Rev. X1, DWGS50467 Rev. X2, | DWGS50467 Rev. 1, DWGS50467 Rev. 2 | DWGS50467 Rev. A, DWGS50467 Rev. B |
| Also referred to as: | “Draft” | “Controlled Draft” | “Released” |
| Point of control: | Controlled by Team -- typically through Share Point and attached to lab books and protocol/reports when needed for design control traceability or IP support | Controlled Document in Document Control | Controlled Document in Document Control |
| Completeness: | In general, will be incomplete (though must be complete enough for design control traceability) | Can be incomplete (though must be complete enough for design control traceability) | Typically complete unless QAP/template specifies completeness in “staged releases” (eg., DFMECAs) |
| Revision tracking: | No formal change tracking required | Formal change tracking required for DMR documents. For most document types, tracking history from numeric revisions is removed when moving to alpha. | Formal change tracking required. |

# Appendix D – Regulatory Requirements

Regulatory requirements for the product are derived from the regulatory bodies in the various target markets for the product.

Listed below are some sources of regulatory requirements that apply to the types of products manufactured by Covidien –

Medical Device Directive (MDD), European Council Directive 93 / 42/ EEC (14 June 1993)

ISO 13485 – Medical devices – Quality management systems – Requirements for regulatory purposes.

CAN/CSA ISO 13485 – Canadian Medical Device Regulations

21 CFR 820 Quality System Regulation for Medical Devices

21 CFR 812 Investigational Device Exemptions

21 CFR 801 Labeling

Medical Devices Regulations, Health Canada, May, 1998, SOR/98-282 (CMDR)

Australian Therapeutic Goods (Medical Device) Regulations 2002

Japan’s Pharmaceutical Affairs Law (PAL)

China Regulation on Supervision and Administration of Medical Devices (Official Draft of Revision of State Council Decree 276)

Reference the Health Industries:  Medical Device Industry website (http://www.ita.doc.gov/td/health/regulations.html) for regulatory requirements for countries not listed above.

While not attempting to be a complete list, this appendix identifies regulatory standards of concern. The Project Core Team must work with Regulatory to determine what is specifically required for a particular project and product.

## Regulatory Submissions

The following deliverables produced by PDP3 are generally required to compile and complete submissions for each of the following global regulatory agencies –

| Deliverable | US |  | EU | Canada | Australia | Japan | China |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | PMA | 510k |  |  |  |  |  |
| Final Labels & Information for Use (IFU) (English Only) | √ | √ | N/A | N/A | N/A | √ | √ |
| Final Labels & Information for Use (IFU) (Translated into required languages) | N/A | N/A | √ | √ | √ | N/A | N/A |
| Design Verification Test (DVT) Protocol | √ | √ | √ | √ | √ | √ | √ |
| Design Validation (DVAL) Test Protocol(s) (including animal testing) | √ | √ | √ | √ | √ | √ | √ |
| Design Verification Test (DVT) Report | √ | √ | √ | √ | √ | √ | √ |
| Design Validation (DVAL) Test Report(s) | √ | √ | √ | √ | √ | √ | √ |
| Accelerated & Real Time Aging Study Protocol (AASP) | √ | √ | √ | √ | √ | √ | √ |
| Accelerated Aging Study Report(s) (AASR) | √ | √ | √ | √ | √ | √ | √ |
| Biocompatibility Validation Protocol | √ | √ | √ | √ | √ | √ | √ |
| Biocompatibility Validation Report (BVR) | √ | √ | √ | √ | √ | √ | √ |
| China Chemical Testing | N/A | N/A | N/A | N/A | N/A | N/A | √ |
| Risk Management Workbook (PHA, UFMECA, & DFMECA) | √ | √ | √ | √ | √ | √ | √ |
| Design Drawings | √ | √ | √ | √ | √ | √ | √ |
| Packaging Validation Protocol (PKP) | √ | √ | √ | √ | √ | √ | √ |
| Packaging Validation Report (PVR) | √ | √ | √ | √ | √ | √ | √ |
| Sterilization Validation Protocol | √ | √ | √ | √ | √ | √ | √ |
| Sterilization Validation Report | √ | √ | √ | √ | √ | √ | √ |
| Bioburden, LAL Bacterial Endotoxin, EO Residual Report | √ | √ | √ | √ | √ | √ | √ |
| Most current annual re-validation of sterilization cycle. | √ | N/A | √ | √ | √ | √ | √ |
| MRI Testing | √ | √ | √ | √ | √ | √ | √ |
| Performance Qualification (PQ) Protocol | √ | N/A | √ | √ | √ | √ | √ |
| Performance Qualification (PQ) Report | √ | N/A | √ | √ | √ | √ | √ |
| Product Process Qualification (PPQ) Protocol | √ | N/A | √ | √ | √ | √ | √ |
| Product Process Qualification (PPQ) Report | √ | N/A | √ | √ | √ | √ | √ |
| Risk Management Workbook (PFMECA) | √ | N/A | √ | √ | √ | √ | √ |
| Risk Management Plan | N/A | N/A | √ | √ | √ | √ | √ |
| Risk Management Report | N/A | N/A | √ | √ | √ | √ | √ |
| Specification Conformance Matrix | √ | √ | √ | √ | √ | √ | √ |
| Quality Monitoring Plan (reference Process Master Validation Matrix (PMVM)) | √ | N/A | √ | √ | √ | √ | √ |
| Clinical Evaluation Report – MDD requirements | N/A | N/A | √ | √ | N/A | √ | √ |
| Supplier certifications for new products. | N/A | N/A | √ | √ | √ | √ | √ |
| List of suppliers and the components they supply – for new products. | √ | N/A | √ | √ | √ | √ | √ |
| Draft of DHF Checklist | √ | √ | √ | √ | √ | √ | √ |
| Detailed list of product materials including MSDS sheets for each material and colorant. For colorants also need: chemical name and CAS number, applicable CFR sections, specific amount of each color additive by weight percent of the colored component and total amount (e.g. µg, ppm) in the device | √ | √ | √ | √ | √ | √ | √ |
| Local product type testing | N/A | N/A | N/A | N/A | N/A | N/A | √ |
| Clinical expert report – TGA requirements | N/A | N/A | N/A | N/A | √ | N/A | N/A |
| Standards conformance forms (details compliance per individual section of the standard) | √ | √ | N/A | √ | N/A | N/A | N/A |
| Clinical study report – IDE requirements | √ | √ | N/A | N/A | N/A | √ | √ |

[This page left intentionally blank]

# Appendix E – Device Master Records

The Device Master Records (DMR) contains all the information necessary to repeatedly produce devices that are within specification.  The deliverables which are part of the DMR do not all have to physically reside in the same location; instead the DMR Index can point to where the information can be found.  The DMR can point to the Design History File (DHF) for information related to the device.  Table 22 provides all required content for the DMR, a possible index for the DMR, and where the information would typically reside.

Table 22 – Device Master Record

| DMR Index | Contents | Location |
| --- | --- | --- |
| Device Specification | - Overall Description - Functional Characteristics - Performance Characteristics | DHF – SCM, EDS |
| Manufacturing Information | - Top assembly drawing - Bill of Materials (including list of all components used in the device) - Procurement documentation - Specifications - Drawings - Certificate of compliance requirements - Supplier Assessment procedures - Device documentation - Assembly drawings - Subassembly drawings - Assembly procedures - Test specifications - Equipment, lines and procedures - Assembly lines - Molds - Machine maintenance procedures - Calibration procedures - Setup procedures - Operating procedures - Process flow charts - Sterilization Procedures - Procedures for ethylene oxide - Handling and flow procedures - Cycle parameter specifications - Diagrams for loading products in the chamber - Production control documentation - Inspection procedures - Test procedures - Blank job travelers - Blank inspection/test forms - Instrument charts - Reporting forms - Approved deviations | DMR – Manufacturing Process Specifications |
| Labeling | - Labeling artwork - Label and IFU drawings - Labeling drawings - Label/labeling review procedures and forms - Production control procedures and history record forms - Instruction for Use - Customer feedback forms | DMR - Manufacturing Process Specifications |
| Packaging | - Package drawings (usually includes labeling information) - Closure drawings - Filling and/or packaging procedures - Packing procedures - Special shipment procedures | DMR - Manufacturing Process Specifications |
| Storage Requirements | - Temperature - Humidity - Shelf Life | DHF – Technical Specifications |
| Inspection Procedures | - Incoming Inspection - In-process Inspections - Finished devices - Process control charts | DMR - Manufacturing Process Specifications |
| Test Procedures | - Incoming test procedures - In-process test procedures - Finished devices - Process control charts | DMR - Manufacturing Process Specifications |
| Final Release | - Release document review list - Distribution procedures - Blank device history record forms | DMR - Manufacturing Process Specifications |

# Appendix F – Deliverable Descriptions

The deliverables created as a result of the Product Development Process are listed in Table 23 below.  The following information is listed for each deliverable –

Phase – the PDP3 phase in which the deliverable is created, approved or updated.

Deliverable ID – a unique identifier for each PDP deliverable.

Deliverable/Review Name – a unique name for each PDP deliverable or review.

Description – a brief description of the information contained in the deliverable.

Key Inputs – other deliverables that contain information that is required to complete the deliverable.

References – additional references that can be used to better understand the deliverable.  Standards, SOPs, templates and auxiliary processes can be used as references.

Reference the Tailoring / RACI Matrix for a detailed description of the deliverable and functional responsibilities for its creation.

Table 23 – Deliverable Definition

| Phase | Del ID | Del/Review Name | Description | Key Inputs |
| --- | --- | --- | --- | --- |
| 3-Dev | 3-004 | Accelerated & Real Time Aging Study Protocol and Accelerated Aging (Interim Time Point) Report - Approved | The Accelerated & Real Time Aging Study Protocol defines testing that subjects samples to elevated temperatures for specific periods of time, to simulate the effects of real-time aging and provides data which allows the manufacturer to accurately predict the effect of real-time aging on his package/product. The Accelerated Aging (Interim Time Point) Report describes the results of testing to confirm earlier prediction based on accelerated aging, related to the effect of aging on the package/product. | • Technical Specifications • Regulatory Plan |
| 4-Qual | 4-024 | Accelerated Aging Study (Final Time Point) Report - Approved | The Accelerated Aging Study (Final Time Point) Report describes the results of testing to confirm earlier prediction based on accelerated aging, related to the effect of aging on the package/product. | • Accelerated Aging Studies Report - Approved (Short Term) • Real Time & Accelerated Aging Study Protocol • Regulatory Plan |
| 4-Qual | 4-008 | Bioburden & LAL Bacterial Endotoxin Report - Approved | Bio-burden / LAL Bacterial Endotoxin Report details the results of testing to determine whether pyrogens are present and the total number of viable microorganisms in or on a medical device, container or component after completion of all in-process steps before sterilization. | • Sterilization Validation Protocol |
| 5-Laun | 5-005 | Additional Accelerated & Real-Time Aging Study Report(s) - Approved | The Additional Accelerated & Real-Time Aging Study Report(s) describe the results of testing to confirm earlier prediction based on accelerated aging and/or real-time aging, related to the effect of aging on the package/product. | • Accelerated & Real Time Aging Study Protocol • Regulatory Plan • Project Plan |
| 3-Dev | 3-005 | Biocompatibility Test Protocol/Report - Approved | The Biocompatibility Test Protocol discusses the method of verifying that materials used in the design and coming into contact with the patient’s blood are biocompatible. The Biocompatibility Test Report describes the results of the biocompatibility testing performed as specified in the Biocompatibility Validation Protocol. | • DMVM • ISO 10993 • Biocompatibility Test Protocol |
| 4-Qual | 4-015 | Clinical Evaluation Report - Approved | The Clinical Evaluation Report is required for CE and international submissions, and is approved early in Qualification. | • DFMECA • CER data • Regulatory Plan • Clinical Study Report |
| 4-Qual | R4-02 | Clinical Readiness Review | The Clinical Readiness Review is a technical review. The purpose of this review is to ascertain that the product is ready for Human Use. |  |
| 2-Feas | 2-001 | Clinical Development Plan (CDP) - Approved | The Clinical Development Plan defines the clinical strategy to support the development and launch of the proposed product. Only completed if clinical study/PPT is being planned. | • Project Plan • Technical Specifications • Regulatory Plan |
| 4-Qual | 4-027 | Clinical Development Plan (CDP) - Updated | The Clinical Development Plan defines the clinical strategy to support the development and launch of the proposed product. | • Clinical Development Plan - Approved • Project Plan • Technical Specifications • Regulatory Plan |
| 4-Qual | 4-021 | Clinical Study Protocol/Report – Approved | The Clinical Study Protocol/Report describes the requirements and results of activities performed in a clinical setting, as defined in the Clinical Development Plan. Only completed if clinical study is planned. | • Clinical Development Plan |
| 5-Laun | 5-002 | Complaint Evaluation Standard (CES) - Approved | The Complaint Evaluation Standard is a template for how to evaluate complaints specific to the product. | • RMW • PFMECA |
| 1-Conc | R1-02 | Concept Selection Technical Review | The Concept Selection Technical Review is a technical review where the core team reviews what has been learned in the Concept Innovation Phase. Most importantly the review covers the concepts considered, how they compared against the Critical to Success Points and other key metrics and documents a rationale for the concept selected. |  |
| 1-Conc | 1-011 | Concept Selection Report - Approved | The Concept Selection Report details the different concepts considered and compares them based on technical risk, medical risk, reliability, manufacturability, IP landscape, and other areas of key importance to the project. | • Testing data from breadboards, models & prototypes • Project Charter • Critical to Quality Points • User Needs • Competitive Product Analysis data • VOC Data |
| 2-Feas | 2-011 | Design Challenge Testing Protocol(s)/Report(s) - Approved | The Design Challenge Testing Protocol(s) are test protocols that challenge the design, to prove that the design is robust. Results are documented in a Design Challenge Test Report. | • Technical Specifications • DESIGN DRAWINGS - Numeric • RMW • Design Characterization Matrix |
| 3-Dev | 3-014 | Design Change Assessment (DCA) - Approved | The Design Change Assessment (DCA) is used to document the impact assessment of all design/label changes. | • DMR • DVT Report • DVAL Report |
| 4-Qual | 4-018 | Design Change Assessment (DCA) - Updated | The Design Change Assessment (DCA) is used to document the impact assessment of all design/label changes. | • DMR • DVT Report • DVAL Report • OQ, PQ, PPQ Reports |
| 2-Feas | 2-003 | Design Characterization Matrix (DCM) – Approved | The Design Characterization Plan (DCP) defines the activities and testing to be performed to develop the engineering design specifications and to define challenge testing to be performed to ensure those specifications are correct. | • Design Characterization Matrix - Draft • Technical Specifications |
| 2-Feas | R2-03 | Design Freeze Technical Review | The Design Freeze Review is a technical review. In this review the design outputs and Pre DVT results are reviewed to ensure that the design is ready to be frozen. |  |
| 1-Conc | 1-013 | Design History File (DHF) – Initiation | The Design History File (DHF) is a compilation of records which describes the design history of a finished device. | • 21 CFR, Part 820 • Project Plan • New and/or updated DHF deliverable references |
| 2-Feas | 2-023 | Design History File (DHF) – Updated | The Design History File (DHF) is a compilation of records which describes the design history of a finished device. | • Project Plan - Updated • New and/or updated DHF deliverable references |
| 3-Dev | 3-017 | Design History File (DHF) – Updated | The Design History File (DHF) is a compilation of records which describes the design history of a finished device. | • Project Plan - Updated • New and/or updated DHF deliverable references |
| 4-Qual | 4-033 | Design History File (DHF) – Updated | The Design History File (DHF) is a compilation of records which describes the design history of a finished device. | • Project Plan - Updated • New and/or updated DHF deliverable references |
| 5-Laun | 5-006 | Design History File (DHF) – Updated | The Design History File (DHF) is a compilation of records which describes the design history of a finished device. | • Project Plan - Updated • New and/or updated DHF deliverable references |
| 2-Feas | 2-009 | Design Test Method Qualification Protocol(s)/Report(s) - Approved | The Design Test Method Qualification Protocol(s) define the testing to be performed to validate test methods to be used in Design Verification and / or Design Validation. Results are documented in Design Test Method Qualification Testing Report(s). | • Design Test Method Document(s) |
| 3-Dev | 3-008 | Design Validation (DVAL) Test Protocol(s)/Report(s) - Approved | Design Validation Test Protocol(s) defines the testing to be performed to validate the design. Results are documented in the Design Validation Test Report(s). | • DVAL Protocol - Draft • DMVM |
| 2-Feas | 2-021 | Design Master Verification / Validation Matrix (DMVM) - Approved | The Design Master Verification / Validation Matrix (DMVM) defines tasks and artifacts required to provide objective evidence that the design complies with the customer and product requirements. | • User Needs • Technical Specifications |
| 4-Qual | 4-011 | Design Master Verification / Validation Matrix (DMVM) - Updated | The Design Master Verification / Validation Matrix (DMVM) defines tasks and artifacts required to provide objective evidence that the design complies with the customer and product requirements. | • DMVM - Approved • OQ, PQ, PPQ Reports • Design TM Qualification Report(s) |
| 1-Conc | 1-009 | Design Test Method (TM) Documents - Draft | The Design Test Method (TM) Document is a compilation of the test methods required to evaluate the competing concepts and competitive products, DVT and the qualification of these methods to ensure that data obtained from them is valid. | • User Needs • Draft Technical Specifications • Draft Design Drawings |
| 2-Feas | 2-008 | Design Test Method (TM) Documents - Approved | The Design Test Method Document is a compilation of the test methods required to evaluate the competing concepts and competitive products, DVT and the qualification of these methods to ensure that data obtained from them is valid. | • Test Methods Document - Draft • User Needs • Technical Specifications • Design Drawings |
| 2-Feas | 2-020 | Design Test Method Capability and Suitability Summary - Approved | The Test Method Capability and Suitability Summary ensures that any new test methods used to perform design and/or process verification or validation testing are suitable for the intended purposes and capable of producing valid results. | • Design Test Method Qualification Report(s) |
| 3-Dev | 3-003 | Design Verification Test (DVT) Protocol/Report - Approved | Design Verification Test (DVT) Protocol defines the testing to be performed (including EO Residuals if applicable) to verify the design and describes the results obtained as a result of performing verification activities defined in the DMVM. Results are documented in the Design Verification Test (DVT) Report. | • DVT Protocol - Draft • DMVM |
| 3-Dev | 3-001 | Device Master Records (DMR) (Numeric Rev) – Approved | The Device Master Records (DMR) is a compilation of records containing an array of documents which define the manufacturing process, procedures and specifications for a finished device. | • Concept Selection Report • Technical Specifications • Design Drawings • Labels & IFU |
| 4-Qual | 4-006 | Device Master Records (DMR) (Alpha Rev) – Updated | The Device Master Records (DMR) is a compilation of records containing an array of documents which define the manufacturing process, procedures and specifications for a finished device. | • DMR – Approved • Concept Selection Report • Technical Specifications • Design Drawings • Labels & IFU |
| 2-Feas | R2-02 | Design for Excellence (DFX) Technical Review | The DFX Review is a technical review. In this review the core team reviews what is known about the product to be built, and what has been learned of the processes needed to build it. |  |
| 1-Conc | 1-008 | Design Drawings (X# Rev) – Draft | Engineering Design Specifications are the array of documents which translate the technical specifications to design outputs from which the physical design can occur. | • Technical Specifications • Design Characterization Matrix - Draft |
| 2-Feas | 2-006 | Design Drawings (Numeric Rev) - Approved | Design Drawings are the array of documents which translate the technical specifications to design outputs from which the physical design can occur. | • Design Drawings – X# Rev • Technical Specifications • Design Characterization Matrix |
| 4-Qual | 4-007 | Design Drawings (Alpha Rev) - Updated | Design Drawings are the array of documents which translate the technical specifications to design outputs from which the physical design can occur. | • Design Drawings – Numeric Rev • Technical Specifications • DVT Report • DVAL Report |
| 2-Feas | 2-017 | Equipment Development Files (EDF) (Initial Requirements) – Approved | The Equipment Development File (EDF) contains all the necessary documentation required to qualify and release for use any equipment, tooling, fixture or other item that is necessary for design control or production. Installation Qualification (IQ) verifies that utility, operational, calibration and maintenance parameters have been considered and developed per manufacturing specifications. | • PMVM • Process Characterization Plan |
| 4-Qual | 4-003 | Equipment Development Files (EDF) – Updated | The Equipment Development File (EDF) contains all the necessary documentation required to qualify and release for use any equipment, tooling, fixture or other item that is necessary for design control or production. Installation Qualification (IQ) verifies that utility, operational, calibration and maintenance parameters have been considered and developed per manufacturing specifications. | • EDF (Initial Requirements) - Approved • PMVM • Process Equipment Qualifications |
| 2-Feas | 2-010 | Feasibility Report - Approved | The Feasibility Report details consolidation of design options through an iterative process of design and testing. Functional and performance data is generated and documented to demonstrate that the consolidated design will, with reasonable confidence of success, meet User Needs. | • Testing data from breadboards, models & prototypes • Project Plan • Critical to Quality Points • User Needs • Competitive Product Analysis data • VOC Data • Concept Selection Report |
| 4-Qual | 4-032 | Field / Physician Training Materials & Records - Approved | The Field / Physician Training Materials define the content of the materials created to train field personnel and physicians. Field / Physician Training Records provide a record of all field personnel and physicians that have been trained to use the product. | • Launch Plan |
| 2-Feas | 2-007 | Global Trade Item Number(s) (GTIN) – Approved | The Global Trade Item Number (GTIN) is a unique 14-digit number used to identify Covidien products. The GTIN is used to encode Covidien's product model numbers in the UCC/EAN-128 bar code format. | • Update LST-003 • Project Plan • SKU/PCN • Package Quantity |
| 1-Conc | 1-012 | IP Planning Session – Approved | Legal group signs that an IP Planning session was held with the project manager to plan out IP milestones and deliverables. | • Project Plan |
| 4-Qual | 4-020 | IP Review Record for First in Human Readiness – Approved | Acknowledgement from Legal providing approval that a project may proceed with human use of the device. | • Regulatory Approval(s) • IP Planning Session |
| 4-Qual | 4-031 | IP Review Record for Launch – Approved | Acknowledgement from Legal providing approval that a project may proceed with product launch. | • Regulatory Approval(s) • IP Planning Session |
| 2-Feas | 2-004 | Labels & Information for Use (IFU) - Draft | Labels and Information for Use (IFU) are the finished labels, any labeling affixed to the device and/or any accompanying literature. | • User Needs • RMW |
| 3-Dev | 3-009 | Labels & Information for Use (IFU) - Approved | Labels and Information for Use (IFU) are the finished labels, any labeling affixed to the device and/or any accompanying literature. | • Labels & IFU - Draft • Budget for IFUs & labels |
| 3-Dev | 3-015 | Launch Plan - Approved | The Launch Plan specifies activities related to launch, including training, Marketing support testing, models | • Project Plan |
| 4-Qual | 4-026 | Launch Plan - Updated | The Launch Plan specifies activities related to launch, including training, Marketing support testing, models | • Launch Plan - Approved |
| 5-Laun | R5-01 | Lessons Learned Review | The purpose of the Lessons Learned Review is to gather the project team to review the entire project from initiation to closure, and to determine what areas of the project and processes worked well and what areas did not. |  |
| 5-Laun | 5-001 | Limited Market Release (LMR) / Market Surveillance | Controlled release of product to the commercial market, utilizing additional quality control over inventory release, following product launch. | • Launch Plan • Regulatory Approval(s) |
| 4-Qual | 4-010 | Manufacturing Transfer Plan - Approved | The Manufacturing Transfer Plan (MTP) identifies the requirements that must be met for successfully transferring the knowledge & responsibility to manufacture the device from Process Development to Manufacturing. | • Supplier Development Report(s) • OQ, PQ, PPQ Report(s) • PFMECA • PMVM • Launch Plan • Project Plan |
| 4-Qual | R4-03 | Manufacturing Transfer Technical Review | The Manufacturing Transfer Readiness Review is a technical review. The purpose of the review is to verify that device design, its components and configuration have been correctly translated into production specifications and is ready to begin transferring responsibility to production & Sustaining Engineering. |  |
| 5-Laun | R5-02 | Manufacturing Transfer Update Review | The Manufacturing Transfer Update Review is a technical review. The purpose of the review is to officially transfer the product to the Support Engineering function. |  |
| 1-Conc | 1-014 | Marketing Claims Matrix | Marketing Claims are the claims marketing desires to make regarding a new product or new feature. | • Project Plan • VOC |
| 4-Qual | 4-028 | Marketing Communications (MARCOMM) – Approved | Marketing Communications is the output of the requirements specified in the Launch Plan for marketing literature. | • Launch Plan |
| 4-Qual | 4-025 | Marketing Support Materials & Claims Testing – Approved | Marketing Support Materials and Claims Testing are the deliverables that R&D provides to support marketing during launch of the product. | • Launch Plan • Marketing Claims Matrix |
| 1-Conc | R1-03 | Phase 1 Gate Review | The Phase 1 Gate Review assesses that the project is worthy of continuation into the Feasibility Phase and that remaining technical, project and product safety risks are manageable. It approves the budget, resources, and proposed schedule to continue with the project. |  |
| 2-Feas | R2-04 | Phase 2 Gate Review | The purpose of the Phase 2 Gate Review is to assess project progress against the plan, review issues and risks that could impact the schedule, cost or resources for the project. |  |
| 3-Dev | R3-02 | Phase 3 Gate Review | The purpose of the Phase 3 Gate Review is to assess project progress against the plan, review issues and risks that could impact the schedule, cost or resources for the project. |  |
| 4-Qual | R4-04 | Phase 4 Gate Review | The main objective of the Phase 4 Gate Review is to determine if the product is ready to enter into full scale production. |  |
| 5-Laun | R5-03 | Phase 5 Gate Review (Project Close-out) | The Phase 5 Gate Review is a non-technical review. This review is used to close out the project. |  |
| 3-Dev | 3-007 | Packaging Validation Protocol/Report - Approved | The Packaging Validation Protocol defines the testing to be performed to validate the packaging that the product is shipped in, both design & process. Results are documented in the Packaging Validation Report. | • Technical Specifications • Design Drawings • PMVM for process • DMVM for verification |
| 4-Qual | 4-022 | Physician Preference Test (PPT) Protocol/Report - Approved | The Physician Preference Test describes the “actual use environment” testing to be performed to assess market acceptance. Results are documented in the Physician Preference Test Report. | • Launch Plan |
| 2-Feas | 2-018 | Pre-Design Verification Test (Pre-DVT) | The Pre-Design Verification Test (Pre-DVT) is the preliminary run on Verification tests to determine if any issues exists prior to entering the Development Phase. | • Technical Specifications • Design Drawings • DVT Protocol – Draft • Design Test Method Qualification Report(s) |
| 4-Qual | 4-004 | Pre-Product Product Qualification (Pre-PPQ) Report - Approved | The Pre-Product Product Qualification (Pre-PPQ) Report documents the results of a preliminary test run of the draft PPQ Protocol to ascertain whether the processes are ready to undergo a successful validation. | • PMVM • DMR • Equipment/Process Qualification Report(s) • Draft PPQ Protocol • OQ, PQ Report(s) • Process Test Method Qualification Report(s) |
| 2-Feas | 2-012 | Process Characterization Plan – Approved | The Process Characterization Plan (PCP) describes activities that will be performed to assess the feasibility and availability of different process technologies and select the appropriate process to support the manufacturing needs of the proposed product. | • Process Characterization Plan – Draft • PMVM • Project Plan |
| 3-Dev | 3-002 | Process Characterization Report – Approved | The Process Characterization Report describes the results of all process characterization activities called out in the Process Characterization Plan. | • Process Characterization Plan |
| 4-Qual | 4-001 | Equipment/Process Qualification Protocols/Reports – Approved | The Equipment/Process Qualification Protocols/Reports define testing required and performed to qualify equipment and/or processes used to manufacture the product. Documented evidence provides a high degree of assurance that a specific process will consistently produce a product meeting its pre-determined specifications and quality characteristics. | • PMVM • Process Characterization Report • IQ’s • PFMECA • DVT Report • Sterilization Validation Report • Packaging Validation Report • DMR • Process Test Method Qualification Report(s) |
| 2-Feas | 2-015 | Process Master Validation Matrix (PMVM) - Draft | The Process Master Validation Matrix describes the strategy, requirements, re-validation requirements and validation results performed to validate the product’s assembly processes and packaging. | • PFMECA • Manufacturing Process Specs • RMW • DMVM |
| 3-Dev | 3-011 | Process Master Validation Matrix (PMVM) - Approved | The Process Master Validation Matrix describes the strategy, requirements, re-validation requirements and validation results performed to validate the product’s assembly processes and packaging. | • PMVM - Draft • PFMECA • Process Characterization Report |
| 4-Qual | 4-012 | Process Master Validation Matrix (PMVM) - Updated | The Process Master Validation Matrix describes the strategy, requirements, re-validation requirements and validation results performed to validate the product’s assembly processes and packaging. | • PMVM – Approved in Development Phase • OQ, PQ, PPQ Reports • Process Equipment Qualification Reports • Process TM Qualifications • DMR |
| 4-Qual | 4-002 | Process Test Method Qualification Protocol/Report – Approved | The Process Test Method Qualification Protocol(s) define the testing to be performed to validate test methods to be used in Process Validation. Results are documented in the Process Test Method Qualification Report. | • PMVM • Test Method Document(s) (MP’s, QP’s) |
| 2-Feas | 2-005 | Product Branding Name (Initiation) – Approved | Submit production information for product name search. | • Project Plan |
| 4-Qual | 4-009 | Product Performance Qualification (PPQ) Protocol/Report - Approved | The Product Performance Qualification Protocol specifies how validation will be conducted, including test parameters, final product characteristics, manufacturing equipment, and decision points on what constitutes acceptable test results. Results are documented in the Product Performance Qualification (PPQ) Report. | • PMVM • Equipment/Process Qualification Report(s) |
| 4-Qual | R4-01 | Product Performance Qualification (PPQ) Readiness Technical Review | The PPQ Readiness Review is a technical review. The purpose of the review is to verify that all documents, information, materials, resources, and equipment needed for process validation are available. |  |
| 1-Conc | 1-001 | Project Charter - Approved | The Project Charter is the document that defines the key parameters of the product, the estimated time and resources required to accomplish this scope, and the key elements of the business case that show this product is a good fit into the overall investment portfolio. | • Business Case • Budget • Resources |
| 1-Conc | 1-002 | Project Plan - Approved | The Project Plan contains a detailed description of project scope, work breakdown, performance measures, and resources | • Project Charter |
| 2-Feas | 2-022 | Project Plan - Updated | The Project Plan contains a detailed description of project scope, work breakdown, performance measures, and resources | • Project Plan – Approved in Concept Phase |
| 3-Dev | 3-016 | Project Plan - Updated | The Project Plan contains a detailed description of project scope, work breakdown, performance measures, and resources | • Project Plan - Updated in Feasibility Phase |
| 4-Qual | 4-030 | Project Plan - Updated | The Project Plan contains a detailed description of project scope, work breakdown, performance measures, and resources | • Project Plan - Updated in Development Phase |
| 1-Conc | 1-005 | Regulatory Plan - Approved | The Regulatory Plan defines the regulatory strategy needed to support the development and launch of the proposed product | • Project Plan |
| 2-Feas | 2-013 | Regulatory Plan - Updated | The Regulatory Plan defines the regulatory strategy needed to support the development and launch of the proposed product | • Regulatory Plan – Approved in Concept Phase |
| 4-Qual | 4-029 | Regulatory Plan - Updated | The Regulatory Plan defines the regulatory strategy needed to support the development and launch of the proposed product | • Regulatory Plan - Updated in Feasibility Phase |
| 5-Laun | 5-003 | Regulatory Plan – Updated | The Regulatory Plan defines the regulatory strategy needed to support the development and launch of the proposed product. | • Launch Plan • Regulatory Plan – Approved in Qualification Phase • Regulatory Approvals |
| 4-Qual | 4-019 | Regulatory Submission(s) & Technical File(s)/ Approval(s) | The Regulatory Submission is the task of submitting the required documentation to regulatory authorities for approval to market the product. | • Regulatory Plan - Approved |
| 4-Qual | 4-017 | Risk Management File Index (RMFI) – Approved | The Risk Management File is a set of records and other documents that are produced by risk management. | • Risk Management Report |
| 1-Conc | 1-003 | Risk Management Plan (RMP) – Approved | The Risk Management Plan defines the risk management activities to be performed throughout the phases of product development. | • Project Plan • ISO 14971 • User Needs |
| 4-Qual | 4-016 | Risk Management Report, pre-Clinical/PPT – Approved | The Risk Management Report documents the results of the risk management activities as defined in the Risk Management Plan for pre-clinical/PPT activities. | • Risk Management Plan • RMW • PFMECA • DVT Report • DVAL Report • OQ, PQ, PPQ Report(s) |
| 4-Qual | 4-023 | Risk Management Report, post-Clinical/PPT – Approved | The Risk Management Report documents the results of the risk management activities as defined in the Risk Management Plan based on results of clinical/PPT activities. | • Risk Management Plan • Clinical Study Report • PPT Report |
| 1-Conc | 1-006 | Risk Management Workbook (RMW) – PHA, UFMECA, DFMECA – Approved | The Risk Management Workbook (RMW) includes the PHA, UFMECA, DFMECA and PFMECA. The Procedure Hazard Analysis (PHA) is a top-down assessment of hazards and clinical effects. The User FMECA and Design FMECA are bottom-up analyses of potential failures of the product to document the effect and identify actions to eliminate or reduce the probability of occurrence of the potential failure. | • Risk Management Plan |
| 2-Feas | 2-019 | Risk Management Workbook (RMW) – PHA, UFMECA, DFMECA – Updated | The Risk Management Workbook (RMW) includes the PHA, UFMECA, DFMECA and PFMECA. The Procedure Hazard Analysis (PHA) is a top-down assessment of hazards and clinical effects. The User FMECA and Design FMECA are bottom-up analyses of potential failures of the product to document the effect and identify actions to eliminate or reduce the probability of occurrence of the potential failure. | • RMW – PHA, UFMECA, DFMECA - Approved |
| 3-Dev | 3-012 | Risk Management Workbook (RMW) – PHA, UFMECA, DFMECA – Updated | The Risk Management Workbook (RMW) includes the PHA, UFMECA, DFMECA and PFMECA. The Procedure Hazard Analysis (PHA) is a top-down assessment of hazards and clinical effects. The User FMECA and Design FMECA are bottom-up analyses of potential failures of the product to document the effect and identify actions to eliminate or reduce the probability of occurrence of the potential failure. | • RMW – PHA, UFMECA, DFMECA – Updated in Feasibility Phase • DVT Report • DVAL Report |
| 4-Qual | 4-013 | Risk Management Workbook (RMW) – PHA, UFMECA, DFMECA – Updated | The Risk Management Workbook (RMW) includes the PHA, UFMECA, DFMECA and PFMECA. The Procedure Hazard Analysis (PHA) is a top-down assessment of hazards and clinical effects. The User FMECA and Design FMECA are bottom-up analyses of potential failures of the product to document the effect and identify actions to eliminate or reduce the probability of occurrence of the potential failure. | • RMW – PHA, UFMECA, DFMECA – Updated in Development Phase • OQ, PQ, PPQ Report(s) |
| 2-Feas | 2-014 | Risk Management Workbook (RMW) – PFMECA – Approved | The Risk Management Workbook (RMW) includes the PHA, UFMECA, DFMECA and PFMECA. The Process Failure Modes, Effects & Criticality Analysis (PFMECA) is used to evaluate potential failures of the process, document the effect of the failure and identify actions to eliminate or reduce the probability of occurrence of the potential failure. | • Process Characterization Plan • Design Drawings (X# Rev) • RMW |
| 3-Dev | 3-010 | Risk Management Workbook (RMW) – PFMECA – Updated | The Risk Management Workbook (RMW) includes the PHA, UFMECA, DFMECA and PFMECA. The Process Failure Modes, Effects & Criticality Analysis (PFMECA) is used to evaluate potential failures of the process, document the effect of the failure and identify actions to eliminate or reduce the probability of occurrence of the potential failure. | • RMW – PFMECA - Approved |
| 4-Qual | 4-014 | Risk Management Workbook (RMW) – PFMECA – Updated | The Risk Management Workbook (RMW) includes the PHA, UFMECA, DFMECA and PFMECA. The Process Failure Modes, Effects & Criticality Analysis (PFMECA) is used to evaluate potential failures of the process, document the effect of the failure and identify actions to eliminate or reduce the probability of occurrence of the potential failure. | • RMW – PFMECA – Updated in Development Phase • OQ, PQ, PPQ Reports |
| 1-Conc | 1-007 | Specification Conformance Matrix (SCM) – User Needs - Draft | The Specification Conformance Matrix (SCM) provides the relationship between user needs, technical specifications, design and testing. | • User Needs • Design Drawings – X# Rev • DVT Protocol • DVAL Protocol |
| 1-Conc | 1-010 | Specification Conformance Matrix (SCM) – Technical Specifications - Draft | The Specification Conformance Matrix (SCM) provides the relationship between user needs, technical specifications, design and testing. | • SCM - User Needs - Draft • Design Drawings – X# Rev • DVT Protocol • DVAL Protocol |
| 2-Feas | 2-016 | Specification Conformance Matrix (SCM) – User Needs & Technical Specifications (Design Inputs) - Approved | The Specification Conformance Matrix (SCM) provides the relationship between user needs, technical specifications, design and testing. | • SCM - Draft • User Needs • Technical Specifications • Design Drawings • DVT Protocol • DVAL Protocol |
| 3-Dev | 3-013 | Specification Conformance Matrix (SCM) - Updated | The Specification Conformance Matrix (SCM) provides the relationship between user needs, technical specifications, design and testing. | • SCM – Approved in Feasibility Phase • Design Drawings • DVT Report • DVAL Report |
| 3-Dev | 3-006 | Sterilization Validation Protocol/Report - Approved | The Sterilization Validation Protocol discusses the method of verifying the sterilization method and dose selected [reduces the bio-burden levels to an acceptable level (SAL)]. The Sterilization Validation Report describes the results of the sterilization studies performed as specified in the Sterilization Protocol. | • Technical Specifications • DMVM • PMVM |
| 5-Laun | 5-004 | Summary Technical Document (STeD) – Approved | Global submission dossier. | • Regulatory Plan • Regulatory Approvals |
| 2-Feas | 2-002 | Supplier Development Plan - Approved | The Supplier Development Plan describes the supply chain strategy for the proposed product that addresses the planning, sourcing, and distribution aspects of the product’s supply chain. | • Project Plan • Concept Selection Report |
| 4-Qual | 4-005 | Supplier Development Report - Approved | The Supplier Development Report describes the results of activities performed to develop and qualify suppliers, as defined in the Supplier Development Plan. | • Supplier Development Plan |
| 2-Feas | R2-01 | Technical Specifications Review | The Technical Specifications Review is a technical review. In this review, the technical specifications are reviewed against the characteristics of good requirements. Additionally, the tracing from technical specifications up to user needs is reviewed to ensure that all user needs have been translated to technical specifications. |  |
| 1-Conc | R1-01 | User Needs Review | The User Needs Review is a technical review of the user needs. In this review the user needs are reviewed for characteristics of good requirements. Additionally, results of any validation of the user needs with the customer are also reviewed at this point. |  |
| 3-Dev | R3-01 | Verification & Validation (V&V) Technical Review | The Verification and Validation Technical Review involves a detailed review of verification and validation data, including design verification, design validation, short term accelerated aging (if complete), biocompatibility, packaging studies, and sterilization validation, China chemical testing |  |
| 1-Conc | 1-004 | Voice of the Customer Plan (VOCP) - Approved | The Voice of the Customer Plan (VOCP) defines VOC activities throughout the product development cycle. | • Project Charter |

# Appendix G – Glossary

| Term | Meaning |
| --- | --- |
| Adoption Justification | An approved file document that leverages existing validation or verification work for use on a new project. Includes justification for why existing validation or verification work is acceptable for use on a new project. |
| baseline | A snapshot in time representing the current agreed-upon, reviewed, and approved set of requirements for a specific product release.[8] |
| bioburden | The population of living microorganisms, bacterial and fungal, on a raw material, component, finished product or packaged. [7] |
| breadboard | A proof-of-concept modeling technique that represents how a product will work, but not how a product will look. |
| Critical to Quality (CTQ) | CTQs (Critical to Quality) are the key measurable characteristics of a product or process whose performance standards or specification limits must be met in order to satisfy the customer. They align improvement or design efforts with user needs. |
| Critical to Success Points (CTS) | Those critical few requirements necessary for commercial success. |
| cross-functional team | A team consisting of representatives from the various functions involved in product development. Typically including Marketing, R&D, Manufacturing, Supply Chain, Regulatory, Clinical, Legal, and Quality. |
| design convergence | Merging of distinct technologies into a unified whole. |
| Design for Manufacturability (DFM) | The systematic consideration of manufacturing issues in the design and development process, facilitating the fabrication of the product’s components and their assembly into the overall product. |
| Design History File (DHF) | The DHF is a compilation of records which describes the design history of a finished device. [2] |
| design input | The physical and performance requirements of a device that are used as a basis for device design. [2] |
| design output | The results of a design effort at each design phase and at the end of the total design effort. The finished design output is the basis for the device master record. The total finished design output consists of the device, its packaging and labeling, and the device master record. [2] |
| design review | A documented, comprehensive, systematic examination of a design to evaluate the adequacy of the design requirements, to evaluate the capability of the design to meet these requirements, and to identify problems. [2] |
| Design Failure Mode, Effects & Criticality Analysis (DFMECA) | The DFMECA is located in the Risk Management Workbook (RMW) and is a bottom-up analysis of potential failures of the product to document the effect and identify actions to eliminate or reduce the probability of occurrence of the potential failure. |
| Device Master Records (DMR) | A compilation of records containing the procedures and specifications for a finished device. [2] |
| essential characteristics | Characteristics that are necessary for the proper functioning of the device. |
| Installation Qualification (IQ) | Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification and that the recommendations of the supplier of the equipment are suitably considered. [4] |
| intellectual property | Information, including proprietary knowledge, technical competencies, and design information, which provides commercially exploitable competitive benefit to an organization. |
| line extension | A derivative product that adds or modifies features without significantly changing the product functionality. |
| Marketing Claims Matrix | The Marketing Claims matrix is a tool used to identify the claims marketing desires to make regarding a new product or new feature. It is used by the team to determine the evidence needed to support the various claims. It also acts as a log to record what claims have been substantiated and which have not/cannot. |
| Operational Qualification (OQ) | Establishing by objective evidence process control limits and action levels which result in product that meets all predetermined requirements. [4] |
| Performance Qualification (PQ) | Establishing by objective evidence that the process, under anticipated conditions, consistently produces a product which meets all predetermined requirements. [4] |
| Product Performance Qualification (PPQ) | Demonstrating that the process has not adversely affected the finished product and that the product meets its predetermined specification and quality attributes. [20] |
| Procedure Hazard Analysis (PHA) | The Procedure Hazard Analysis (PHA) is located in the Risk Management Workbook (RMW) and is a top-down risk analysis tool used in the overall risk management of a project. The PHA is designed to be performed early in the product development process to provide input into the DFMECA and to provide preliminary recommendations for risk mitigation. |
| process | A set of interrelated or interacting activities which transforms inputs into outputs. ISO] |
| process validation | Establishing by objective evidence that a process consistently produces a result or product meeting its predetermined requirements. [4] |
| prototype | A physical model of the new product concept. Depending upon the purpose, prototypes may be non-working, or functionally working. |
| reliability | The characteristic of a device, or any component thereof, expressed as a probability that it will perform its required functions under defined conditions for a specified operation period. [FDA 90-423] |
| robustness | Robustness is defined as the degree to which a system operates correctly in the presence of exceptional inputs or stressful environmental conditions. [IEEE] |
| software validation | Confirmation by examination and provision of objective evidence that software specifications conform to user needs and intended uses, and that the particular requirements implemented through software can be consistently fulfilled. [19] |
| User Failure Modes, Effects & Criticality Analysis (UFMECA) | The User FMECA is located in the Risk Management Workbook (RMW) and is a sequential risk analysis tool used to determine the risk level associated with foreseeable misuse of the product. |
| user needs | Problems that the user intends to solve with the purchase of a good or service. Also known as “customer needs”. |
| validation | To provide objective evidence that device specifications conform with user needs and intended uses, and to provide objective evidence that devices produced in accordance with the design satisfy user needs and intended uses. [2] |
| verification | To assure that design inputs are adequate, and to confirm that design outputs meet the design input requirements. [2] |
| Voice of the Customer (VOC) | VOC is a market research technique that produces a detailed set of customer wants and needs, organized into a hierarchical structure, and then prioritized in terms of relative importance and satisfaction with current alternatives. |

# Appendix H – References

General references used in development of the PDP3 Process Guide are listed below.  Footnotes throughout the document designate additional references related to specific content.

| [1] | The Toyota Product Development System, Morgan, James M and Liker, Jeffrey K., New York, NY: Productivity Press, 2006. |
| --- | --- |
| [2] | Design Control Guidance For Medical Device Manufacturers, FDA, March 11, 1997. |
| [3] | Six Sigma for Dummies, Gygi, Craig, DeCarlo, Neil, and Williams, Bruce, Indianapolis, Indiana, Wiley Publishing, Inc., 2005. |
| [4] | Quality Management Systems – Process Validation Guidance, GHTF, Edition 2, January 2004. |
| [5] | ISO 14971 Medical devices – Application of risk management to medical devices, Third edition, 2012. |
| [6] | ISO 13485 Medical devices – Quality management systems – Requirements for regulatory purposes, Second edition, 2003. |
| [7] | The Biomedical Quality Auditor Handbook, Milwaukee, WI: ASQ Quality Press, 2003. |
| [8] | Software Requirements, Wiegers, Karl E., Redmond, WA: Microsoft Press, 2003. |
| [9] | Measurement Systems Analysis Reference Manual, DaimlerChrysler Corporation, Ford Motor Company Corporation, General Motors Corporation, Third Edition, March 2002. |
| [10] | CMMI for Development Version 1.2, Carnegie Mellon University, Software Engineering Institute, August 2006. |
| [11] | Interpretation of FDA Quality System Regulations (QSR), Medical Device School, Inc., 2006. |
| [12] | Systems Engineering Handbook, International Council on Systems Engineering (INCOSE), 2010. |
| [13] | Quality Management System – Medical Devices – Guidance on the Control of Products and Services Obtained from Suppliers, GHTF 2008. |
| [14] | Lean Product and Process Development, Ward, Allen C., Cambridge, MA: The Lean Enterprise Institute, Inc., 2007. |
| [15] | www.isixsigma.com, I Six Sigma organization. |
| [16] | www.npd-solutions.com, New Product Development Solutions organization. |
| [17] | Assessing Biocompatibility, SGS, 2007. |
| [18] | Preparing for Successful Design Transfer, Zenner, Terry, Medical Device & Diagnostic Industry, 2001. |
| [19] | General Principles of Software Validation; Final Guidance for Industry and FDA Staff, FDA, January 11, 2002. |
| [20] | Quality System Manual, Process Validation, FDA, December 1996. |

[THIS PAGE INTENTIONALLY LEFT BLANK]

For more on detail on the major deliverables of the PDP, see Appendix F.

“Design verification protocols” refers to all protocols used for design verification, including the Sterilization Validation and Accelerated Aging protocols.

Refer to the section on Verifying the Product Design for a description of Accelerated Aging and Real Time Aging studies.

For more information on Project Close Out, see the S2insight Project Management Methodology 2.0, June 2009.

This is commonly referred to as the “5 Whys”.

Reference [1] The Toyota Product Development System. New York, NY: Productivity Press.  (2006)

On occasion more than one concept is selected to continue to the next phase of development.  This can be done to minimize risk in the Feasibility Phase by having another alternative or to continue to converge the concepts early in the Feasibility Phase.

Product requirements are not usually prioritized.

For references on statistical techniques use QAP 20.2 and www.itl.nist.gov/div898/handbook/.

Reference http://www.itl.nist.gov/div898/handbook/ for more information on Control Charts.

ISO standards exist for the most common methods of sterilization, including the following requirements for development, validation, and routine control of sterilization processes for medical devices -

- ISO 11135 - Sterilization of health care products - ethylene oxide.

- ISO 17665 - Sterilization of health care products - moist heat.

- ISO 11137 - Sterilization of health care products - radiation.

The number 10-6 denotes the number of products that will have a living microorganism following the sterilization process.

For more information on design for robustness, go to the section entitled Developing a Robust Design in Appendix A6.

For references on sample sizes and confidence intervals, use QAC 20.2.

These definitions of waste are from The Toyota Product Development System by James M. Morgan and Jeffrey K. Liker.

PDP3 does not address the mechanics of how design changes are performed.  Please reference QAP 5.1 Document Control System for requirements on requesting and approving changes.

For biocompatibility, sterilization, packaging, and Accelerated Aging studies if not performing testing, adoption document is needed.

For biocompatibility, sterilization, packaging, and Accelerated Aging studies if not performing testing, adoption document is needed.

For more information on the DMR refer to http://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/PostmarketRequirements/QualitySystemsRegulations/MedicalDeviceQualitySystemsManual/